## output
60Gy 8F | Mar 05 2018 | Test | Difference | Nov 22 2017 | Test | Difference | Nov 17 2017 | Test | Difference | 48Gy4F 60Gy5F | Feb 14 2018 | Test | Difference | Feb 27 2018 | Test | Difference | Sept 29 2017 | Test | Difference | Aug 29 2017 | Test | Difference | 54Gy 3F | SABR1 | Test | Difference | SABR2 | Test | Difference | SABR3 | Test | Difference | SABR4 | Test | Difference | SABR5 | Test | Difference | SABR6 | Test | Difference
Patient: | 0 | AA, BB |  | 0 | AA, BB |  | 0 | AA, BB |  | Patient: | 0 | AA, BB |  | 0 | AA, BB |  | 0 | AA, BB |  | 0 | AA, BB |  | Patient: | PT1 | PT1 | TRUE | PT2 | PT2 | TRUE | PT3 | PT3 | TRUE | PT4 | PT4 | TRUE | PT5 | PT5 | TRUE | PT6 | PT6 | TRUE
CR#: | 0 | 11 |  | 0 | 11 |  | 0 | 11 |  | CR#: | 0 | 11 |  | 0 | 11 |  | 0 | 11 |  | 0 | 11 |  | CR#: | 1 | 1 | 0 | 2 | 2 | 0 | 3 | 3 | 0 | 4 | 4 | 0 | 5 | 5 | 0 | 6 | 6 | 0
Site: | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  | Site: | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  | Site: | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE
Plan Name: | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | Plan Name: | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNR | LUNR | TRUE | Plan Name: | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE
Prescription Dose (cGy): | 0 | 6000 |  | 0 | 6000 |  | 0 | 6000 |  | Prescription Dose (cGy): | 4800 | 4800 | 0 | 4800 | 4800 | 0 | 4800 | 4800 | 0 | 4800 | 4800 | 0 | Prescription Dose (cGy): | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0
Fractions: | 0 | 8 |  | 0 | 8 |  | 0 | 8 |  | Fractions: | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  | Fractions: | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0
GTV Volume (cc) | 5.1 | 5.1 | 0 | 3.9 | 3.9 | 0 | 12.2 | 12.2 | 0 | GTV Volume (cc) | 4.7 | 4.7 | 0 | 7.3 | 7.3 | 0 | 8 | 8 | 0 | 8.1 | 8.1 | 0 | GTV Volume (cc) | 1.33 | 1.3 | 0.03 | 3.5 | 3.5 | 0 | 2.09 | 2.1 | -0.01 | 1.3 | 1.3 | 0 | 1.84 | 2.5 | -0.66 |  | ??
ITV Volume (cc): | 5.1 | 5.1 | 0 | 3.9 | 3.9 | 0 | 12.1 | 12.1 | 0 | ITV Volume (cc): | 4.7 | 4.7 | 0 | 7.3 | 7.3 | 0 | 8 | 8 | 0 | 8.1 | 8.1 | 0 | ITV Volume (cc): | 1.35 | 1.3 | 0.05 | 3.5 | 3.5 | 0 | 2.08 | 2.1 | -0.02 | 1.3 | 1.3 | 0 | 2.47 | 2.5 | -0.03 | 11.94 | 11.9 | 0.04
PTV Volume (cc) | 20.3 | 20.3 | 0 | 17.2 | 17.2 | 0 | 36.8 | 36.8 | 0 | PTV Volume (cc) | 17.4 | 17.4 | 0 | 23.8 | 23.8 | 0 | 26.1 | 26.1 | 0 | 24.2 | 24.2 | 0 | PTV Volume (cc) | 7.35 | 7.3 | 0.05 | 14.7 | 14.7 | 0 | 10.08 | 10.1 | -0.02 | 8.2 | 8.2 | 0 | 10.94 | 10.9 | 0.04 | 34.52 | 34.5 | 0.02
Total Lung Volume (cc) | 2229.1 | 2229.1 | 0 | 3687.7 | 3687.7 | 0 | 3431.8 | 3431.8 | 0 | Total Lung Volume (cc) | 3365.2 | 3365.2 | 0 | 4403.6 | 4403.6 | 0 | 4489.8 | 4489.8 | 0 | 2030 | 2030 | 0 | Total Lung Volume (cc) | 4689.47 | 4689.5 | -0.03 | 6102.7 | 6102.7 | 0 | 3151.03 | 3151 | 0.03 | 2926.04 | 2926 | 0.04 | 5232 | 5232.3 | -0.3 | 7331 | 7331 | 0
Plan Normalization Value (%) | 0.891 | ?? |  | 0.909 | ?? |  | 0.86 | ?? |  | Plan Normalization Value (%) | ?? | ?? | TRUE | 0.85 | ?? |  | 0.894 | ?? |  | 0.822 | ?? |  | Plan Normalization Value (%) | 66.3% | ?? |  | 87.7% | ?? |  | 72.4% | ?? |  | 82.3% | ?? |  | 86.7% | ?? |  | 71.9% | ??
Dose @COM-PTV (cGy) | 111.7% | ?? |  | 114.7% | ?? |  | 122.2% | ?? |  | Dose @COM-PTV (%) | ?? | ?? | TRUE | 126.1% | ?? |  | 119.4% | ?? |  | 122.3% | ?? |  | Dose @COM-PTV (cGy) | 6385.6 | ?? |  | 6520.1 | ?? |  | 6547.2 | ?? |  | 6389 | ?? |  | 7203.6 | ?? |  | 6711.3 | ??
PTV- Minimum Dose | 84.7% | 84.7% | 0% | 90.3% | 90.3% | 0% | 85.9% | 85.9% | 0% | PTV- Minimum Dose (%) | 90.4% | 90.4% | 0% | 91% | 91% | 0% | 89.5% | 89.5% | 0% | 87% | 87% | 0% | PTV- Minimum Dose | 90.9% | 90.92% | -0.01852% | 88.84% | 88.84% | 0% | 90.65% | 90.65% | 0% | 91.02% | 99.12% | -8.106% | 90.37% | 90.66% | -0.2889% | 84.48% | 84.48% | 0.003704%
PTV - V100(%) | 95% | 95% | 0% | 95% | 95% | 0% | 95% | 95% | 0% | PTV - V100(%) | 95% | 95% | 0% | 95% | 95% | 0% | 95% | 95% | 0% | 95% | 95% | 0% | PTV - V100(%) | 95% | 95.59% | -0.5867% | 95% | 95.05% | -0.04558% | 95% | 94.82% | 0.1793% | 95% | 100.6% | -5.583% | 95% | 95.36% | -0.3642% | 95% | 95.06% | -0.05594%
PTV - V90 (%) | 99.83% | 99.83% | 0% | 100% | 100% | 0% | 99.92% | 99.92% | 0% | PTV - V90 (%) | 100% | 100% | 0% | 100% | 100% | 0% | 100% | 100% | 0% | 99.94% | 99.94% | 0% | PTV - V90 (%) | 100% | 100.6% | -0.6177% | 100% | 100% | -0.0449% | 99.9% | 99.81% | 0.08911% | 100% | 100.6% | -0.5884% | 100% | 100.4% | -0.3835% | 99% | 99.73% | -0.729%
Location V105% - PTV (cc) = | 0.1 | 0.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Location V105% - PTV (cc) = | 0 | 0 | 0 | 0.1 | 0.1 | 0 | 0 | 0 | 0 | 0.1 | 0.1 | 0 | Location V105% - PTV (cc) = | 0.05 | 0 | 0.05 | 0.1 | 0 | 0.1 | 0.05 | 0 | 0.05 | 0 | ?? |  | 0.03 | 0 | 0.03 | 2.35 | 2.4 | -0.05
Volume V100% (cc) = | 21.7 | 21.7 | 0 | 17.6 | 17.6 | 0 | 38.9 | 38.9 | 0 | Volume V100% (cc) = | 17.2 | 17.2 | 0 | 25.4 | 25.4 | 0 | 26 | 26 | 0 | 25 | 25 | 0 | Volume V100% (cc) = | 7.86 | 7.9 | -0.04 | 15.4 | 15.4 | 0 | 11.61 | 11.6 | 0.01 | 14.88 | 14.9 | -0.02 | 11.6 | 11.6 | 0 | 40.34 | 40.3 | 0.04
Location D³2cm (%) = | 55.2 | 55.2 | 0 | 49.7 | 49.7 | 0 | 54 | 54 | 0 | Location D³2cm (%) = | 46.1 | 46.1 | 0 | 61.5 | 61.5 | 0 | 51.1 | 51.1 | 0 | 58.6 | 58.6 | 0 | Location D³2cm (%) = | 48.6 | 48.613 | -0.012963 | 53.8 | 53.837 | -0.037037 | 52.8704 | 52.8704 | 0 | 50.2204 | 50.2204 | 0 | 51.3148 | 51.313 | 0.0018519 | 67.1296 | 67.1204 | 0.0092593
Volume V50% (cc) = | 89.3 | 89.3 | 0 | 79.1 | 79.1 | 0 | 161.5 | 161.5 | 0 | Volume | 71.1 | 71.1 | 0 | 105.5 | 105.5 | 0 | 112.2 | 112.2 | 0 | 98.4 | 98.4 | 0 | Volume(R50) | 38.89 | 38.9 | -0.01 | 70.61 | 70.6 | 0.01 | 57.78 | 57.8 | -0.02 | 69.86 | 69.9 | -0.04 | 52.9 | 52.9 | 0 | 152.4 | 152.4 | 0
Mean Dose (contralateral lung) | 102 | 102 | 1.4211e-14 | 72 | 72 | 0 | 120 | 120 | 0 | Mean Dose (contralateral lung) | 67.2 | 67.2 | 0 | 62.4 | 62.4 | 0 | 72 | 72 | 0 | 120 | 120 | 0 | Mean Dose (contralateral lung) | 52.8 | 52.8 | 0 | 50.1 | 50.7 | -0.6 | 71.8 | 71.8 | 0 | 92 | 92 | 0 | 79.1 | 79.1 | 0 | 84.7 | 84.7 | 0
Mean Dose (Total lung) | 300 | 300 | 0 | 162 | 162 | 2.8422e-14 | 468 | 468 | 0 | Mean Dose (Total lung) | 288 | 288 | 0 | 201.6 | 201.6 | 0 | 259.2 | 259.2 | 0 | 321.6 | 321.6 | 0 | Mean Dose (Total lung) | 173.7 | 173.7 | 0 | 185.9 | 185.9 | 0 | 262.6 | 262.6 | 0 | 286.6 | 286.6 | 0 | 205.4 | 205.4 | 0 | 233.7 | 233.7 | 0
V20 (Total Lung) | 0.037159 | 0.037095 | 6.4067e-05 | 0.021945 | 0.021916 | 2.9133e-05 | 0.055226 | 0.055139 | 8.6833e-05 | V20 (Total Lung) in % | 0.028168 | 0.028086 | 8.1467e-05 | 0.015966 | 0.015924 | 4.16e-05 | 0.025632 | 0.025566 | 6.6667e-05 | 0.035633 | 0.035542 | 9.1133e-05 | V20 (Total Lung) in % | 0.015 | 0.014668 | 0.0003317 | 0.021 | 0.020838 | 0.00016217 | 0.0254 | 0.025313 | 8.6512e-05 | 0.035 | 0.034963 | 3.6569e-05 | 0.018 | 0.018387 | -0.00038746 | 0.022 | 0.02265 | -0.00065025
Aorta (max point dose) | 1572 | 1572 | 0 | 558 | 558 | 1.1369e-13 | 1812 | 1812 | 0 | Lung-Basic Function | 213.6357 | 212.8316 | 0.80406 | 174.1611 | 173.4656 | 0.69548 | 252.7461 | 251.8933 | 0.85276 | 151.1593 | 150.7393 | 0.41994 | Lung-Basic Function V11.6Gy= | 161 | 161.522 | -0.522 | 247.68 | 247.678 | 0.002 | 160.4 | 160.899 | -0.499 | 199 | 198.413 | 0.587 | 245 | 245.359 | -0.359 | 355 | 354.187 | 0.813
Aorta V60Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lung-Pneumonitis | 194.4702 | 194.1227 | 0.34751 | 157.246 | 156.9356 | 0.31031 | 231.9408 | 231.5637 | 0.37714 | 141.1597 | 140.9765 | 0.18324 | Lung-Pneumonitis V12.4Gy= | 146 | 146.462 | -0.462 | 232.3 | 232.3 | 0 | 149 | 149.052 | -0.052 | 187 | 186.755 | 0.245 | 222 | 221.778 | 0.222 | 327 | 327.509 | -0.509
Artery-Pulmonary (max point dose) | 36 | 36 | 0 | 36 | 36 | 0 | 1134 | 1134 | -2.2737e-13 | Aorta (max point dose) | 547.2 | 547.2 | 0 | 979.2 | 979.2 | 0 | 734.4 | 734.4 | -1.1369e-13 | 1104 | 1104 | 0 | Aorta (max point dose) | 675.6 | 675.6 | 0 | 1340.7 | 1340.7 | 0 | 959.6 | 959.6 | 0 | 1518.1 | 1518.1 | 0 | 1154 | 1154.1 | -0.1 | 1560 | 1559.5 | 0.5
Artery-Pulmonary V60Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Aorta V40Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Artery-Pulmnory (max point dose) | 606.8 | 606.8 | 0 | 220.9 | 220.9 | 0 | 887.1 | 887.1 | 0 | 33.2 | 33.2 | 0 | 867 | 867.1 | -0.1 | 1529 | 1529.1 | -0.1
Spinal Canal (max point dose) | 1086 | 1086 | 2.2737e-13 | 1644 | 1644 | -2.2737e-13 | 906 | 906 | 0 | Artery-Pulmonary (max point dose) | 96 | 96 | 0 | 1046.4 | 1046.4 | 0 | 619.2 | 619.2 | 0 | 806.4 | 806.4 | 0 | Spinal Canal (max point dose) | 681.8 | 681.8 | 0 | 842.6 | 842.6 | 0 | 570.4 | 570.4 | 0 | 907 | 907 | 0 | 1723 | 1722.9 | 0.1 | 473 | 473.3 | -0.3
Ipsilat. Brach. Plex. (max point dose) | ?? | ?? | TRUE | 2304 | 2304 | 0 | ?? | ?? | TRUE | Artery-Pulmonary V40Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Spinal Canal-PRV 5mm (max point dose) | 722 | 722 | 0 | 924 | 924 | 0 | 607.3 | 607.3 | 0 | 1013.3 | 1013.3 | 0 | 1843 | 1843.4 | -0.4 | 501 | 500.8 | 0.2
Ipsilat. Brach. Plex. V30Gy= | ?? | ?? | TRUE | 0 | 0 | 0 | ?? | ?? | TRUE | Spinal Canal (max point dose) | 1012.8 | 1012.8 | 0 | 600 | 600 | 0 | 518.4 | 518.4 | 0 | 1123.2 | 1123.2 | 0 | Ipsilat Brach.Plex (max point dose) | ?? | ?? | TRUE | 0 | ?? |  | 0 | ?? |  | ?? | ?? | TRUE | ?? | ?? | TRUE | 0 | ??
Heart (max point dose) | 1932 | 1932 | -2.2737e-13 | 18 | 18 | 0 | 1038 | 1038 | 0 | Spinal Canal V20.8Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Skin V30Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Heart V60Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Spinal Canal V13.6Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Heart (max point dose) | 851.5 | 851.5 | 0 | 67.3 | 67.3 | 0 | 969.1 | 969.1 | 0 | 1799.9 | 1799.9 | 0 | 175 | 174.9 | 0.1 | 173 | 173.2 | -0.2
Esophagus (max point dose) | 1032 | 1032 | 0 | 1278 | 1278 | 0 | 942 | 942 | 0 | Spinal Canal-PRV 5mm (max point dose) | 1166.4 | 1166.4 | -2.2737e-13 | 648 | 648 | 0 | 556.8 | 556.8 | 0 | 1291.2 | 1291.2 | 0 | Esophagus (max point dose) | 691.2 | 691.2 | 0 | 800 | 800 | 0 | 602.4 | 602.4 | 0 | 761.6 | 761.6 | 0 | 1000 | 999.7 | 0.3 | 878 | 878.1 | -0.1
Chestwall (rib) (max point dose) | 7260 | 7260 | 0 | 7044 | 7044 | 9.0949e-13 | 7320 | 7320 | 0 | Ipsilat. Brach. Plex. (max point dose) | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | *Chestwall (rib) (max point dose) | 2605.1 | 2605.1 | 0 | 2472.2 | 2472.2 | 0 | 4716.9 | 4716.9 | 0 | 3488.3 | 3488.3 | 0 | 2450 | 2449.5 | 0.5 | 6296 | 6296.4 | -0.4
Chestwall (rib) V50Gy= | 13.1625 | 13.149 | 0.013476 | 7.4308 | 7.4222 | 0.0085671 | 8.674 | 8.6619 | 0.012172 | Ipsilat. Brach. Plex. V23.6Gy= | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | *Chestwall (rib) V28.2Gy= | OK | 0 |  | OK | 0 |  | 6.8 | 6.8456 | -0.04564 | 2.5 | 2.549 | -0.04905 | 0 | 0 | 0 | 30 | 29.7168 | 0.2832
Trachea (max point dose) | 0 | 18 |  | 0 | 1452 |  | 0 | 54 |  | Skin V30Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | *Chestwall (rib) V30Gy= | OK | 0 |  | OK | 0 |  | 4.8 | 4.8319 | -0.0319 | 1.2 | 1.2179 | -0.01786 | 0 | 0 | 0 | 24 | 24.5536 | -0.5536
Trachea V60Gy = | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Heart (max point dose) | 1243.2 | 1243.2 | 2.2737e-13 | 1425.6 | 1425.6 | 0 | 96 | 96 | 0 | 129.6 | 129.6 | 0 | Trachea | 0 | 10.4 |  | 0 | 692.2 |  | 0 | 29.7 |  | 0 | 19 |  | 0 | 684.7 |  | 0 | 166.5
Bronchus (max point dose) | 0 | 48 |  | 0 | 96 |  | 0 | 2316 |  | Heart V28Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Bronchus | 0 | 1194 |  | 0 | 443.3 |  | 0 | 1124.7 |  | 0 | 65.1 |  | 0 | 1852.6 |  | 0 | 1504.6
Bronchus V36Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Esophagus (max point dose) | 763.2 | 763.2 | 0 | 1425.6 | 1425.6 | 0 | 566.4 | 566.4 | 0 | 1406.4 | 1406.4 | -2.2737e-13 | Proximal Trachea and Bronchial Tree (max point dose) | 1194 | 1194 | 0 | 692.2 | 692.2 | 0 | 1124.7 | 1124.7 | 0 | 65 | 65.1 | -0.1 | 1853 | 1852.6 | 0.4 | 1504.6 | 1504.6 | 0
Proximal Trachea & Bronch. Tree: (max point dose) | 48 | 48 | 0 | 1452 | 1452 | 0 | 2316 | 2316 | 0 | Esophagus V18.8Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dosimetrist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
Proximal Trachea & Bronch. Tree: V60Gy = | 0 | 0 | 0 | 0 | ?? |  | 0 | ?? |  | *Chestwall (rib) (max point dose) | 2342.4 | 2342.4 | 4.5475e-13 | 5947.2 | 5947.2 | 0 | 4838.4 | 4838.4 | 0 | 5995.2 | 5995.2 | 0 | Physicist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
Stomach and Intestines (max point dose) | ?? | ?? | TRUE | 0 | ?? |  | ?? | ?? | TRUE | *Chestwall (rib) V40Gy= | 0 | 0 | 0 | 10.8067 | 10.7934 | 0.013312 | 0.77112 | 0.76777 | 0.0033482 | 8.9172 | 8.9064 | 0.010805 | Radiation Oncologist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
Stomach and Intestines V36Gy= | ?? | ?? | TRUE | 0 | ?? |  | ?? | ?? | TRUE | *Chestwall (rib) V30Gy= | 0 | 0 | 0 | 23.716 | 23.716 | 0 | 6.5726 | 6.5726 | 0 | 20.3863 | 20.3863 | 0 | NOTES: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dosimetrist: | ?? | ?? | TRUE | 0 | ?? |  | ?? | ?? | TRUE | Trachea (max point dose) | 0 | 19.2 |  | 0 | 19.2 |  | 0 | 628.8 |  | 0 | 1444.8 |  |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Physicist: | ?? | ?? | TRUE | 0 | ?? |  | ?? | ?? | TRUE | Trachea V15.6Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Radiation Oncologist: | ?? | ?? | TRUE | 0 | ?? |  | ?? | ?? | TRUE | Bronchus (max point dose) | 0 | 936 |  | 0 | 585.6 |  | 0 | 849.6 |  | 0 | 1444.8 |  |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
NOTES: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Bronchus V15.6Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | Proximal Trachea and Bronchial Tree (max point dose) | 936 | 936 | 0 | 585.6 | 585.6 | 1.1369e-13 | 849.6 | 849.6 | 1.1369e-13 | 1444.8 | 1444.8 | -2.2737e-13 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | Proximal Trachea and Bronchial Tree V15.6Gy= | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  | 0 | ?? |  |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | Stomach and Intestines (max point dose) | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | Stomach and Intestines V21Gy= | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | Dosimetrist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | Physicist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | 0 | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[10 empty] | NOTES: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[13 empty] | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[13 empty] | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
[13 empty] | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0

## Format Template
60Gy 8F | Mar 05 2018 | Test | Difference | Nov 22 2017 | Test | Difference | Nov 17 2017 | Test | Difference | 48Gy4F 60Gy5F | Feb 14 2018 | Test | Difference | Feb 27 2018 | Test | Difference | Sept 29 2017 | Test | Difference | Aug 29 2017 | Test | Difference | 54Gy 3F | SABR1 | Test | Difference | SABR2 | Test | Difference | SABR3 | Test | Difference | SABR4 | Test | Difference | SABR5 | Test | Difference | SABR6 | Test | Difference
EvaluationSheet 60Gy 8F | DVH Data Mar 05 2018.dvh | Test_results | Difference | DVH Data Nov 22 2017.dvh | Test_results | Difference | DVH Data Nov 17 2017.dvh | Test_results | Difference | EvaluationSheet 48Gy4F 60Gy5F | DVH Data Feb 14 2018.dvh | Test_results | Difference | DVH Data Feb 27 2018.dvh | Test_results | Difference | DVH_Data_Sept_29_2017.dvh | Test_results | Difference | DVH_Data_Aug_29_2017.dvh | Test_results | Difference | EvaluationSheet 54Gy 3F | SABR1.dvh | Test_results | Difference | SABR2.dvh | Test_results | Difference | SABR3.dvh | Test_results | Difference | SABR4.dvh | Test_results | Difference | SABR5.dvh | Test_results | Difference | SABR6.dvh | Test_results | Difference
Patient: | 0 | AA, BB | FALSE | 0 | AA, BB | FALSE | 0 | AA, BB | FALSE | Patient: | 0 | AA, BB | FALSE | 0 | AA, BB | FALSE | 0 | AA, BB | FALSE | 0 | AA, BB | FALSE | Patient: | PT1 | PT1 | TRUE | PT2 | PT2 | TRUE | PT3 | PT3 | TRUE | PT4 | PT4 | TRUE | PT5 | PT5 | TRUE | PT6 | PT6 | TRUE
CR#: | 0 | 11 | -11 | 0 | 11 | -11 | 0 | 11 | -11 | CR#: | 0 | 11 | -11 | 0 | 11 | -11 | 0 | 11 | -11 | 0 | 11 | -11 | CR#: | 1 | 1 | 0 | 2 | 2 | 0 | 3 | 3 | 0 | 4 | 4 | 0 | 5 | 5 | 0 | 6 | 6 | 0
Site: | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | Site: | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | Site: | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE | LUNG | LUNG | TRUE
Plan Name: | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | Plan Name: | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNR | LUNR | TRUE | Plan Name: | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE | LUNR | LUNR | TRUE | LUNL | LUNL | TRUE
Prescription Dose (cGy): | 0 | 6000 | -6000 | 0 | 6000 | -6000 | 0 | 6000 | -6000 | Prescription Dose (cGy): | 4800 | 4800 | 0 | 4800 | 4800 | 0 | 4800 | 4800 | 0 | 4800 | 4800 | 0 | Prescription Dose (cGy): | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0 | 5400 | 5400 | 0
Fractions: | 0 | 8 | -8 | 0 | 8 | -8 | 0 | 8 | -8 | Fractions: | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | Fractions: | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0 | 3 | 3 | 0
GTV Volume (cc) | 5.1 | 5.1 | 0 | 3.9 | 3.9 | 0 | 12.2 | 12.2 | 0 | GTV Volume (cc) | 4.7 | 4.7 | 0 | 7.3 | 7.3 | 0 | 8 | 8 | 0 | 8.1 | 8.1 | 0 | GTV Volume (cc) | 1.33 | 1.3 | 0.03 | 3.5 | 3.5 | 0 | 2.09 | 2.1 | -0.01 | 1.3 | 1.3 | 0 | 1.84 | 2.5 | -0.66 |  | ?? | FALSE
ITV Volume (cc): | 5.1 | 5.1 | 0 | 3.9 | 3.9 | 0 | 12.1 | 12.1 | 0 | ITV Volume (cc): | 4.7 | 4.7 | 0 | 7.3 | 7.3 | 0 | 8 | 8 | 0 | 8.1 | 8.1 | 0 | ITV Volume (cc): | 1.35 | 1.3 | 0.05 | 3.5 | 3.5 | 0 | 2.08 | 2.1 | -0.02 | 1.3 | 1.3 | 0 | 2.47 | 2.5 | -0.03 | 11.94 | 11.9 | 0.04
PTV Volume (cc) | 20.3 | 20.3 | 0 | 17.2 | 17.2 | 0 | 36.8 | 36.8 | 0 | PTV Volume (cc) | 17.4 | 17.4 | 0 | 23.8 | 23.8 | 0 | 26.1 | 26.1 | 0 | 24.2 | 24.2 | 0 | PTV Volume (cc) | 7.35 | 7.3 | 0.05 | 14.7 | 14.7 | 0 | 10.08 | 10.1 | -0.02 | 8.2 | 8.2 | 0 | 10.94 | 10.9 | 0.04 | 34.52 | 34.5 | 0.02
Total Lung Volume (cc) | 2229.1 | 2229.1 | 0 | 3687.7 | 3687.7 | 0 | 3431.8 | 3431.8 | 0 | Total Lung Volume (cc) | 3365.2 | 3365.2 | 0 | 4403.6 | 4403.6 | 0 | 4489.8 | 4489.8 | 0 | 2030 | 2030 | 0 | Total Lung Volume (cc) | 4689.47 | 4689.5 | -0.03 | 6102.7 | 6102.7 | 0 | 3151.03 | 3151 | 0.03 | 2926.04 | 2926 | 0.04 | 5232 | 5232.3 | -0.3 | 7331 | 7331 | 0
Plan Normalization Value (%) | 0.891 | ?? | FALSE | 0.909 | ?? | FALSE | 0.86 | ?? | FALSE | Plan Normalization Value (%) | ?? | ?? | TRUE | 0.85 | ?? | FALSE | 0.894 | ?? | FALSE | 0.822 | ?? | FALSE | Plan Normalization Value (%) | 66.3% | ?? | FALSE | 87.7% | ?? | FALSE | 72.4% | ?? | FALSE | 82.3% | ?? | FALSE | 86.7% | ?? | FALSE | 71.9% | ?? | FALSE
Dose @COM-PTV (cGy) | 111.7% | ?? | FALSE | 114.7% | ?? | FALSE | 122.2% | ?? | FALSE | Dose @COM-PTV (%) | ?? | ?? | TRUE | 126.1% | ?? | FALSE | 119.4% | ?? | FALSE | 122.3% | ?? | FALSE | Dose @COM-PTV (cGy) | 6385.6 | ?? | FALSE | 6520.1 | ?? | FALSE | 6547.2 | ?? | FALSE | 6389 | ?? | FALSE | 7203.6 | ?? | FALSE | 6711.3 | ?? | FALSE
PTV- Minimum Dose | 84.7% | 84.7% | 0% | 90.3% | 90.3% | 0% | 85.9% | 85.9% | 0% | PTV- Minimum Dose (%) | 90.4% | 1573% | -1483% | 91% | 2166% | -2075% | 89.5% | 2336% | -2246% | 87% | 2105% | -2018% | PTV- Minimum Dose | 90.9% | 90.92% | -0.01852% | 88.84% | 88.84% | 0% | 90.65% | 90.65% | 0% | 91.02% | 99.12% | -8.106% | 90.37% | 90.66% | -0.2889% | 84.48% | 84.48% | 0.003704%
PTV - V100(%) | 95% | 95% | 0% | 95% | 95% | 0% | 95% | 95% | 0% | PTV - V100(%) | 95% | 95% | 0% | 95% | 95% | 0% | 95% | 95% | 0% | 95% | 95% | 0% | PTV - V100(%) | 95% | 95.59% | -0.5867% | 95% | 95.05% | -0.04558% | 95% | 94.82% | 0.1793% | 95% | 100.6% | -5.583% | 95% | 95.36% | -0.3642% | 95% | 95.06% | -0.05594%
PTV - V90 (%) | 99.83% | 99.83% | 0% | 100% | 100% | 0% | 99.92% | 99.92% | 0% | PTV - V90 (%) | 100% | 100% | 0% | 100% | 100% | 0% | 100% | 100% | 0% | 99.94% | 99.94% | 0% | PTV - V90 (%) | 100% | 100.6% | -0.6177% | 100% | 100% | -0.0449% | 99.9% | 99.81% | 0.08911% | 100% | 100.6% | -0.5884% | 100% | 100.4% | -0.3835% | 99% | 99.73% | -0.729%
Location V105% - PTV (cc) = | 0.1 | 0.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Location V105% - PTV (cc) = | 0 | 0 | 0 | 0.1 | 0.1 | 0 | 0 | 0 | 0 | 0.1 | 0.1 | 0 | Location V105% - PTV (cc) = | 0.05 | 0 | 0.05 | 0.1 | 0 | 0.1 | 0.05 | 0 | 0.05 | 0 | ?? | FALSE | 0.03 | 0 | 0.03 | 2.35 | 2.4 | -0.05
Volume V100% (cc) = | 21.7 | 21.7 | 0 | 17.6 | 17.6 | 0 | 38.9 | 38.9 | 0 | Volume V100% (cc) = | 17.2 | 17.2 | 0 | 25.4 | 25.4 | 0 | 26 | 26 | 0 | 25 | 25 | 0 | Volume V100% (cc) = | 7.86 | 7.9 | -0.04 | 15.4 | 15.4 | 0 | 11.61 | 11.6 | 0.01 | 14.88 | 14.9 | -0.02 | 11.6 | 11.6 | 0 | 40.34 | 40.3 | 0.04
Location D³2cm (%) = | 55.2 | 55.2 | 0 | 49.7 | 49.7 | 0 | 54 | 54 | 0 | Location D³2cm (%) = | 46.1 | 46.1 | 0 | 61.5 | 61.5 | 0 | 51.1 | 51.1 | 0 | 58.6 | 58.6 | 0 | Location D³2cm (%) = | 48.6 | 48.613 | -0.012963 | 53.8 | 53.837 | -0.037037 | 52.8704 | 52.8704 | 0 | 50.2204 | 50.2204 | 0 | 51.3148 | 51.313 | 0.0018519 | 67.1296 | 67.1204 | 0.0092593
Volume V50% (cc) = | 89.3 | 89.3 | 0 | 79.1 | 79.1 | 0 | 161.5 | 161.5 | 0 | Volume | 71.1 | 71.1 | 0 | 105.5 | 105.5 | 0 | 112.2 | 112.2 | 0 | 98.4 | 98.4 | 0 | Volume(R50) | 38.89 | 38.9 | -0.01 | 70.61 | 70.6 | 0.01 | 57.78 | 57.8 | -0.02 | 69.86 | 69.9 | -0.04 | 52.9 | 52.9 | 0 | 152.4 | 152.4 | 0
Mean Dose (contralateral lung) | 102 | 102 | 1.4211e-14 | 72 | 72 | 0 | 120 | 120 | 0 | Mean Dose (contralateral lung) | 67.2 | 67.2 | 0 | 62.4 | 62.4 | 0 | 72 | 72 | 0 | 120 | 120 | 0 | Mean Dose (contralateral lung) | 52.8 | 52.8 | 0 | 50.1 | 50.7 | -0.6 | 71.8 | 71.8 | 0 | 92 | 92 | 0 | 79.1 | 79.1 | 0 | 84.7 | 84.7 | 0
Mean Dose (Total lung) | 300 | 300 | 0 | 162 | 162 | 2.8422e-14 | 468 | 468 | 0 | Mean Dose (Total lung) | 288 | 288 | 0 | 201.6 | 201.6 | 0 | 259.2 | 259.2 | 0 | 321.6 | 321.6 | 0 | Mean Dose (Total lung) | 173.7 | 173.7 | 0 | 185.9 | 185.9 | 0 | 262.6 | 262.6 | 0 | 286.6 | 286.6 | 0 | 205.4 | 205.4 | 0 | 233.7 | 233.7 | 0
V20 (Total Lung) | 0.037159 | 0.037095 | 6.4067e-05 | 0.021945 | 0.021916 | 2.9133e-05 | 0.055226 | 0.055139 | 8.6833e-05 | V20 (Total Lung) in % | 0.028168 | 0.028086 | 8.1467e-05 | 0.015966 | 0.015924 | 4.16e-05 | 0.025632 | 0.025566 | 6.6667e-05 | 0.035633 | 0.035542 | 9.1133e-05 | V20 (Total Lung) in % | 0.015 | 0.014668 | 0.0003317 | 0.021 | 0.020838 | 0.00016217 | 0.0254 | 0.025313 | 8.6512e-05 | 0.035 | 0.034963 | 3.6569e-05 | 0.018 | 0.018387 | -0.00038746 | 0.022 | 0.02265 | -0.00065025
Aorta (max point dose) | 1572 | 1572 | 0 | 558 | 558 | 1.1369e-13 | 1812 | 1812 | 0 | Lung-Basic Function | 213.6357 | 212.8316 | 0.80406 | 174.1611 | 173.4656 | 0.69548 | 252.7461 | 251.8933 | 0.85276 | 151.1593 | 150.7393 | 0.41994 | Lung-Basic Function V11.6Gy= | 161 | 161.522 | -0.522 | 247.68 | 247.678 | 0.002 | 160.4 | 160.899 | -0.499 | 199 | 198.413 | 0.587 | 245 | 245.359 | -0.359 | 355 | 354.187 | 0.813
Aorta V60Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lung-Pneumonitis | 194.4702 | 194.1227 | 0.34751 | 157.246 | 156.9356 | 0.31031 | 231.9408 | 231.5637 | 0.37714 | 141.1597 | 140.9765 | 0.18324 | Lung-Pneumonitis V12.4Gy= | 146 | 146.462 | -0.462 | 232.3 | 232.3 | 0 | 149 | 149.052 | -0.052 | 187 | 186.755 | 0.245 | 222 | 221.778 | 0.222 | 327 | 327.509 | -0.509
Artery-Pulmonary (max point dose) | 36 | 36 | 0 | 36 | 36 | 0 | 1134 | 1134 | -2.2737e-13 | Aorta (max point dose) | 547.2 | 547.2 | 0 | 979.2 | 979.2 | 0 | 734.4 | 734.4 | -1.1369e-13 | 1104 | 1104 | 0 | Aorta (max point dose) | 675.6 | 675.6 | 0 | 1340.7 | 1340.7 | 0 | 959.6 | 959.6 | 0 | 1518.1 | 1518.1 | 0 | 1154 | 1154.1 | -0.1 | 1560 | 1559.5 | 0.5
Artery-Pulmonary V60Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Aorta V40Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Artery-Pulmnory (max point dose) | 606.8 | 606.8 | 0 | 220.9 | 220.9 | 0 | 887.1 | 887.1 | 0 | 33.2 | 33.2 | 0 | 867 | 867.1 | -0.1 | 1529 | 1529.1 | -0.1
Spinal Canal (max point dose) | 1086 | 1086 | 2.2737e-13 | 1644 | 1644 | -2.2737e-13 | 906 | 906 | 0 | Artery-Pulmonary (max point dose) | 96 | 96 | 0 | 1046.4 | 1046.4 | 0 | 619.2 | 619.2 | 0 | 806.4 | 806.4 | 0 | Spinal Canal (max point dose) | 681.8 | 681.8 | 0 | 842.6 | 842.6 | 0 | 570.4 | 570.4 | 0 | 907 | 907 | 0 | 1723 | 1722.9 | 0.1 | 473 | 473.3 | -0.3
Ipsilat. Brach. Plex. (max point dose) | ?? | ?? | TRUE | 2304 | 2304 | 0 | ?? | ?? | TRUE | Artery-Pulmonary V40Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Spinal Canal-PRV 5mm (max point dose) | 722 | 722 | 0 | 924 | 924 | 0 | 607.3 | 607.3 | 0 | 1013.3 | 1013.3 | 0 | 1843 | 1843.4 | -0.4 | 501 | 500.8 | 0.2
Ipsilat. Brach. Plex. V30Gy= | ?? | ?? | TRUE | 0 | 4.416 | -4.416 | ?? | ?? | TRUE | Spinal Canal (max point dose) | 1012.8 | 1012.8 | 0 | 600 | 600 | 0 | 518.4 | 518.4 | 0 | 1123.2 | 1123.2 | 0 | Ipsilat Brach.Plex (max point dose) | ?? | ?? | TRUE | 0 | ?? | FALSE | 0 | ?? | FALSE | ?? | ?? | TRUE | ?? | ?? | TRUE | 0 | ?? | FALSE
Heart (max point dose) | 1932 | 1932 | -2.2737e-13 | 18 | 18 | 0 | 1038 | 1038 | 0 | Spinal Canal V20.8Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Skin V30Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Heart V60Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Spinal Canal V13.6Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Heart (max point dose) | 851.5 | 851.5 | 0 | 67.3 | 67.3 | 0 | 969.1 | 969.1 | 0 | 1799.9 | 1799.9 | 0 | 175 | 174.9 | 0.1 | 173 | 173.2 | -0.2
Esophagus (max point dose) | 1032 | 1032 | 0 | 1278 | 1278 | 0 | 942 | 942 | 0 | Spinal Canal-PRV 5mm (max point dose) | 1166.4 | 1166.4 | -2.2737e-13 | 648 | 648 | 0 | 556.8 | 556.8 | 0 | 1291.2 | 1291.2 | 0 | Esophagus (max point dose) | 691.2 | 691.2 | 0 | 800 | 800 | 0 | 602.4 | 602.4 | 0 | 761.6 | 761.6 | 0 | 1000 | 999.7 | 0.3 | 878 | 878.1 | -0.1
Chestwall (rib) (max point dose) | 7260 | 7260 | 0 | 7044 | 7044 | 9.0949e-13 | 7320 | 7320 | 0 | Ipsilat. Brach. Plex. (max point dose) | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | *Chestwall (rib) (max point dose) | 2605.1 | 2605.1 | 0 | 2472.2 | 2472.2 | 0 | 4716.9 | 4716.9 | 0 | 3488.3 | 3488.3 | 0 | 2450 | 2449.5 | 0.5 | 6296 | 6296.4 | -0.4
Chestwall (rib) V50Gy= | 13.1625 | 13.149 | 0.013476 | 7.4308 | 7.4222 | 0.0085671 | 8.674 | 8.6619 | 0.012172 | Ipsilat. Brach. Plex. V23.6Gy= | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | *Chestwall (rib) V28.2Gy= | OK | 0 | FALSE | OK | 0 | FALSE | 6.8 | 6.8456 | -0.04564 | 2.5 | 2.549 | -0.04905 | 0 | 0 | 0 | 30 | 29.7168 | 0.2832
Trachea (max point dose) | 0 | 18 | -18 | 0 | 1452 | -1452 | 0 | 54 | -54 | Skin V30Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | *Chestwall (rib) V30Gy= | OK | 0 | FALSE | OK | 0 | FALSE | 4.8 | 4.8319 | -0.0319 | 1.2 | 1.2179 | -0.01786 | 0 | 0 | 0 | 24 | 24.5536 | -0.5536
Trachea V60Gy = | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Heart (max point dose) | 1243.2 | 1243.2 | 2.2737e-13 | 1425.6 | 1425.6 | 0 | 96 | 96 | 0 | 129.6 | 129.6 | 0 | Trachea | 0 | 10.4 | -10.4 | 0 | 692.2 | -692.2 | 0 | 29.7 | -29.7 | 0 | 19 | -19 | 0 | 684.7 | -684.7 | 0 | 166.5 | -166.5
Bronchus (max point dose) | 0 | 48 | -48 | 0 | 96 | -96 | 0 | 2316 | -2316 | Heart V28Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Bronchus | 0 | 1194 | -1194 | 0 | 443.3 | -443.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bronchus V36Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Esophagus (max point dose) | 763.2 | 763.2 | 0 | 1425.6 | 1425.6 | 0 | 566.4 | 566.4 | 0 | 1406.4 | 1406.4 | -2.2737e-13 | Proximal Trachea and Bronchial Tree (max point dose) | 1194 | 1194 | 0 | 692.2 | 692.2 | 0 | 1124.7 | 29.7 | 1095 | 65 | 19 | 46 | 1853 | 684.7 | 1168.3 | 167 | 166.5 | 0.5
Proximal Trachea & Bronch. Tree: (max point dose) | 240 | 48 | 192 | 1452 | 1452 | 0 | 5724 | 2316 | 3408 | Esophagus V18.8Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dosimetrist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
Proximal Trachea & Bronch. Tree: V60Gy = | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | *Chestwall (rib) (max point dose) | 2342.4 | 2342.4 | 4.5475e-13 | 5947.2 | 5947.2 | 0 | 4838.4 | 4838.4 | 0 | 5995.2 | 5995.2 | 0 | Physicist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
Stomach and Intestines (max point dose) | ?? | ?? | TRUE | 0 | ?? | FALSE | ?? | ?? | TRUE | *Chestwall (rib) V40Gy= | 0 | 0 | 0 | 10.8067 | 10.7934 | 0.013312 | 0.77112 | 0.76777 | 0.0033482 | 8.9172 | 8.9064 | 0.010805 | Radiation Oncologist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
Stomach and Intestines V36Gy= | ?? | ?? | TRUE | 0 | ?? | FALSE | ?? | ?? | TRUE | *Chestwall (rib) V30Gy= | 0 | 0 | 0 | 23.716 | 23.716 | 0 | 6.5726 | 6.5726 | 0 | 20.3863 | 20.3863 | 0 | NOTES: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dosimetrist: | ?? | ?? | TRUE | 0 | ?? | FALSE | ?? | ?? | TRUE | Trachea (max point dose) | 0 | 19.2 | -19.2 | 0 | 19.2 | -19.2 | 0 | 628.8 | -628.8 | 0 | 1444.8 | -1444.8
Physicist: | ?? | ?? | TRUE | 0 | ?? | FALSE | ?? | ?? | TRUE | Trachea V15.6Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Radiation Oncologist: | ?? | ?? | TRUE | 0 | ?? | FALSE | ?? | ?? | TRUE | Bronchus (max point dose) | 0 | 936 | -936 | 0 | 585.6 | -585.6 | 0 | 849.6 | -849.6 | 0 | 1444.8 | -1444.8
NOTES: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Bronchus V15.6Gy= | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[10 empty] | Proximal Trachea and Bronchial Tree (max point dose) | 5208 | 936 | 4272 | 1281.6 | 585.6 | 696 | 1526.4 | 849.6 | 676.8 | 1444.8 | 1444.8 | -2.2737e-13
[10 empty] | Proximal Trachea and Bronchial Tree V15.6Gy= | 11.1879 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE | 0 | ?? | FALSE
[10 empty] | Stomach and Intestines (max point dose) | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
[10 empty] | Stomach and Intestines V21Gy= | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
[10 empty] | Dosimetrist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
[10 empty] | Physicist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
[10 empty] | Radiation Oncologist: | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE | ?? | ?? | TRUE
[10 empty] | NOTES: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Chestwall
Type | Data Source | Item | Reference | Test | Difference
60Gy 8F | Mar 05 2018 | Chestwall (rib) V50Gy= | 13.1625 | 13.149 | 0.013476
60Gy 8F | Nov 22 2017 | Chestwall (rib) V50Gy= | 7.4308 | 7.4222 | 0.0085671
60Gy 8F | Nov 17 2017 | Chestwall (rib) V50Gy= | 8.674 | 8.6619 | 0.012172
48Gy4F 60Gy5F | Feb 14 2018 | *Chestwall (rib) (max point dose) | 2342.4 | 2342.4 | 4.5475e-13
48Gy4F 60Gy5F | Feb 14 2018 | *Chestwall (rib) V40Gy= | 0 | 0 | 0
48Gy4F 60Gy5F | Feb 14 2018 | *Chestwall (rib) V30Gy= | 0 | 0 | 0
48Gy4F 60Gy5F | Feb 27 2018 | *Chestwall (rib) (max point dose) | 5947.2 | 5947.2 | 0
48Gy4F 60Gy5F | Feb 27 2018 | *Chestwall (rib) V40Gy= | 10.8067 | 10.7934 | 0.013312
48Gy4F 60Gy5F | Feb 27 2018 | *Chestwall (rib) V30Gy= | 23.716 | 23.716 | 0
48Gy4F 60Gy5F | Sept 29 2017 | *Chestwall (rib) (max point dose) | 4838.4 | 4838.4 | 0
48Gy4F 60Gy5F | Sept 29 2017 | *Chestwall (rib) V40Gy= | 0.77112 | 0.76777 | 0.0033482
48Gy4F 60Gy5F | Sept 29 2017 | *Chestwall (rib) V30Gy= | 6.5726 | 6.5726 | 0
48Gy4F 60Gy5F | Aug 29 2017 | *Chestwall (rib) (max point dose) | 5995.2 | 5995.2 | 0
48Gy4F 60Gy5F | Aug 29 2017 | *Chestwall (rib) V40Gy= | 8.9172 | 8.9064 | 0.010805
48Gy4F 60Gy5F | Aug 29 2017 | *Chestwall (rib) V30Gy= | 20.3863 | 20.3863 | 0
54Gy 3F | SABR1 | *Chestwall (rib) V28.2Gy= | OK | 0 | FALSE
54Gy 3F | SABR1 | *Chestwall (rib) V30Gy= | OK | 0 | FALSE
54Gy 3F | SABR2 | *Chestwall (rib) V28.2Gy= | OK | 0 | FALSE
54Gy 3F | SABR2 | *Chestwall (rib) V30Gy= | OK | 0 | FALSE
54Gy 3F | SABR3 | *Chestwall (rib) V28.2Gy= | 6.8 | 6.8456 | -0.04564
54Gy 3F | SABR3 | *Chestwall (rib) V30Gy= | 4.8 | 4.8319 | -0.0319
54Gy 3F | SABR4 | *Chestwall (rib) V28.2Gy= | 2.5 | 2.549 | -0.04905
54Gy 3F | SABR4 | *Chestwall (rib) V30Gy= | 1.2 | 1.2179 | -0.01786
54Gy 3F | SABR5 | *Chestwall (rib) V28.2Gy= | 0 | 0 | 0
54Gy 3F | SABR5 | *Chestwall (rib) V30Gy= | 0 | 0 | 0
54Gy 3F | SABR6 | *Chestwall (rib) V28.2Gy= | 30 | 29.7168 | 0.2832
54Gy 3F | SABR6 | *Chestwall (rib) V30Gy= | 24 | 24.5536 | -0.5536