## points
FieldDose | FieldRefPointEffectiveDepth | FieldRefPointPointDepth | FieldRefPointSSD | RefPointId | field_name | RefPointId | RefPointX | RefPointY | RefPointZ
97.8 | 12.1 | 12 | 97.4 | Thick d=12, -5 | 20x20 6MV 0Nd | Thick d=12, -5 | -7.5 | 9 | 0
81.3 | 3 | 3 | 97.5 | Thick d=3, -5 | 20x20 6MV 0Nd | Thick d=3, -5 | -7.5 | 0 | 0
124.2 | 6 | 6 | 97.5 | Thick d=6, -5 | 20x20 6MV 0Nd | Thick d=6, -5 | -7.5 | 3 | 0
112.7 | 9 | 9 | 97.5 | Thick d=9, -5 | 20x20 6MV 0Nd | Thick d=9, -5 | -7.5 | 6 | 0
48.1 | - | - | - | Thin d=0, 5 | 20x20 6MV 0Nd | Thin d=0, 5 | 2.5 | 0 | 0
154.4 | 3 | 3 | 100 | Thin d=3, 5 | 20x20 6MV 0Nd | Thin d=3, 5 | 2.5 | 3 | 0
136.1 | 6 | 6 | 100 | Thin d=6, 5 | 20x20 6MV 0Nd | Thin d=6, 5 | 2.5 | 6 | 0
118.4 | 9 | 9 | 100 | Thin d=9, 5 | 20x20 6MV 0Nd | Thin d=9, 5 | 2.5 | 9 | 0
62.1 | 12.1 | 12 | 94.4 | Thick d=12, -5 | 20x20 6MV 3Nd
15.9 | 3 | 3 | 94.5 | Thick d=3, -5 | 20x20 6MV 3Nd
46.7 | 6.1 | 6 | 94.4 | Thick d=6, -5 | 20x20 6MV 3Nd
68.5 | 9 | 9 | 94.4 | Thick d=9, -5 | 20x20 6MV 3Nd
32.2 | - | - | - | Thin d=0, 5 | 20x20 6MV 3Nd
101.2 | 3 | 3 | 97 | Thin d=3, 5 | 20x20 6MV 3Nd
88.9 | 6 | 6 | 97 | Thin d=6, 5 | 20x20 6MV 3Nd
77.2 | 9 | 9 | 97 | Thin d=9, 5 | 20x20 6MV 3Nd
65.4 | 12 | 12 | 91.4 | Thick d=12, -5 | 20x20 6MV 6Nd
10.6 | 3 | 3 | 91.5 | Thick d=3, -5 | 20x20 6MV 6Nd
19 | 6 | 6 | 91.5 | Thick d=6, -5 | 20x20 6MV 6Nd
46.1 | 9 | 9 | 91.4 | Thick d=9, -5 | 20x20 6MV 6Nd
37.6 | - | - | - | Thin d=0, 5 | 20x20 6MV 6Nd
115.2 | 3 | 3 | 94 | Thin d=3, 5 | 20x20 6MV 6Nd
101 | 6 | 6 | 94 | Thin d=6, 5 | 20x20 6MV 6Nd
87.5 | 9 | 9 | 94 | Thin d=9, 5 | 20x20 6MV 6Nd
45.8 | 12.1 | 12.1 | 88.4 | Thick d=12, -5 | 20x20 6MV 9Nd
9.1 | 3 | 3 | 88.5 | Thick d=3, -5 | 20x20 6MV 9Nd
14.4 | 6 | 6 | 88.5 | Thick d=6, -5 | 20x20 6MV 9Nd
21.3 | 9 | 9 | 88.5 | Thick d=9, -5 | 20x20 6MV 9Nd
44.1 | - | - | - | Thin d=0, 5 | 20x20 6MV 9Nd
133.2 | 3 | 3 | 91 | Thin d=3, 5 | 20x20 6MV 9Nd
116.5 | 6 | 6 | 91 | Thin d=6, 5 | 20x20 6MV 9Nd
100.6 | 9 | 9 | 91 | Thin d=9, 5 | 20x20 6MV 9Nd
271.2 |  |  |  | Thick d=12, -5 | Total
116.9 |  |  |  | Thick d=3, -5 | Total
204.3 |  |  |  | Thick d=6, -5 | Total
248.6 |  |  |  | Thick d=9, -5 | Total
162 |  |  |  | Thin d=0, 5 | Total
504 |  |  |  | Thin d=3, 5 | Total
442.5 |  |  |  | Thin d=6, 5 | Total
383.7 |  |  |  | Thin d=9, 5 | Total

## fields
Course | PlanId | FieldId | FieldName | FieldMachineId | FieldEnergyMode | FieldGantryAngle | FieldCollimatorAngle | FieldTableAngle | X1 | X2 | Y1 | Y2 | FieldIsocentreX | FieldIsocentreY | FieldIsocentreZ | FieldSAD | FieldSSD | FieldActualSSD | FieldMonitorUnits | FieldWeightFactor | FieldRefDose | FieldNormFactor | FieldNormMethod | FieldDoseRate | FieldTechnique | FieldIMRT | FieldCalculationWarning | FieldCalculationInfo
AirGap2 | AG30N6XF20 | 20x20 6MV 0Nd | Thin 0 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 0 | 0 | 100 | 100 | - | 157 | 1.634 | 161.7 | 100 | FIELD_ISOC_NORM | 600 MU | STATIC |  | Dose in isocenter is too small | Service: AAA.FieldVolumeDose.10.0.28.9901x64 Servant: p2792@varfaspr17 Calculated Wednesday, October 02, 2019 10:03:19 Client workstation CTX7-ARIA-PV4/5396 Client version 13.6.32.20182 Client time 2019-10-02T10:03:19-04:00 (ISO8601) Anisotropic Analytical Algorithm (Version 10.0.28) There is a new version of this algorithm available (13.6.23). Using DCF protocol version 0.3 AAACalculationOptions.CalculationGridSizeInCM = 0.25 AAACalculationOptions.FieldNormalizationType = 100% to isocenter AAACalculationOptions.HeterogeneityCorrection = ON AAACalculationOptions.StdArcCalculationSegments = 10 AAACalculationOptions.VMATFluenceResolution = High Machine directory: \\VARFASPR13\DCF$\client\BeamData\AAA_10025_Gold... Treatment unit: TR1, energy: 6X Accessory OPENFIELD matches code 00 Input Fluence Pixel Size 0.3125x0.3125 mm2. Field dose normalized 100% to max dose at isocenter plane
AirGap2 | AG30N6XF20 | 20x20 6MV 3Nd | Thin 3 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 3.02 | 0 | 100 | 97 | - | 97 | 1.011 | 106.1 | 100 | FIELD_ISOC_NORM | 600 MU | STATIC |  |  | Service: AAA.FieldVolumeDose.10.0.28.9901x64 Servant: p3900@varfaspr18 Calculated Wednesday, October 02, 2019 10:03:19 Client workstation CTX7-ARIA-PV4/5396 Client version 13.6.32.20182 Client time 2019-10-02T10:03:19-04:00 (ISO8601) Anisotropic Analytical Algorithm (Version 10.0.28) There is a new version of this algorithm available (13.6.23). Using DCF protocol version 0.3 AAACalculationOptions.CalculationGridSizeInCM = 0.25 AAACalculationOptions.FieldNormalizationType = 100% to isocenter AAACalculationOptions.HeterogeneityCorrection = ON AAACalculationOptions.StdArcCalculationSegments = 10 AAACalculationOptions.VMATFluenceResolution = High Machine directory: \\VARFASPR13\DCF$\client\BeamData\AAA_10025_Gold... Treatment unit: TR1, energy: 6X Accessory OPENFIELD matches code 00 Input Fluence Pixel Size 0.3125x0.3125 mm2. Field dose normalized 100% to isocenter
AirGap2 | AG30N6XF20 | 20x20 6MV 6Nd | Thin 6 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 6.02 | 0 | 100 | 94 | - | 104 | 1.009 | 120.9 | 100 | FIELD_ISOC_NORM | 600 MU | STATIC |  |  | Service: AAA.FieldVolumeDose.10.0.28.9901x64 Servant: p3944@varfaspr19 Calculated Wednesday, October 02, 2019 10:03:19 Client workstation CTX7-ARIA-PV4/5396 Client version 13.6.32.20182 Client time 2019-10-02T10:03:19-04:00 (ISO8601) Anisotropic Analytical Algorithm (Version 10.0.28) There is a new version of this algorithm available (13.6.23). Using DCF protocol version 0.3 AAACalculationOptions.CalculationGridSizeInCM = 0.25 AAACalculationOptions.FieldNormalizationType = 100% to isocenter AAACalculationOptions.HeterogeneityCorrection = ON AAACalculationOptions.StdArcCalculationSegments = 10 AAACalculationOptions.VMATFluenceResolution = High Machine directory: \\VARFASPR13\DCF$\client\BeamData\AAA_10025_Gold... Treatment unit: TR1, energy: 6X Accessory OPENFIELD matches code 00 Input Fluence Pixel Size 0.3125x0.3125 mm2. Field dose normalized 100% to isocenter
AirGap2 | AG30N6XF20 | 20x20 6MV 9Nd | Thin 9 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 9.02 | 0 | 100 | 91 | - | 113 | 1.005 | 140.1 | 100 | FIELD_ISOC_NORM | 600 MU | STATIC |  |  | Service: AAA.FieldVolumeDose.10.0.28.9901x64 Servant: p5552@varfaspr13 Calculated Wednesday, October 02, 2019 10:03:19 Client workstation CTX7-ARIA-PV4/5396 Client version 13.6.32.20182 Client time 2019-10-02T10:03:19-04:00 (ISO8601) Anisotropic Analytical Algorithm (Version 10.0.28) There is a new version of this algorithm available (13.6.23). Using DCF protocol version 0.3 AAACalculationOptions.CalculationGridSizeInCM = 0.25 AAACalculationOptions.FieldNormalizationType = 100% to isocenter AAACalculationOptions.HeterogeneityCorrection = ON AAACalculationOptions.StdArcCalculationSegments = 10 AAACalculationOptions.VMATFluenceResolution = High Machine directory: \\VARFASPR13\DCF$\client\BeamData\AAA_10025_Gold... Treatment unit: TR1, energy: 6X Accessory OPENFIELD matches code 00 Input Fluence Pixel Size 0.3125x0.3125 mm2. Field dose normalized 100% to isocenter

## plan
PatientLastName | Virt Sim
PatientFirstName | Test Plans
PatientId | PYVirtSimTest
PatientSex | Unknown
UserId | gsal
IsPlanModified | -
IsAllMUperGyValid
Invalid If condition
CourseId | AirGap2
CourseIntent | QA
PlanId | AG30N6XF20
PlanName | AirGap 3
ImageId | AirGapBase
ImageName | AirGap3
ImageSeriesId | AirGap
ImageModality | CT
ImagingDeviceId | GE Multislice CT
ImagingDeviceModel | LightSpeed RT16
PatientImagingOrientation | Head First-Supine
PatientTreatmentOrientation | Head First-Supine
PlanningCoordinateSystemName | Standard
ImageUserOrigin | User origin DICOM offset =
ApprovalStatus | Unapproved
PlanCreator | gsal
PlanModifier | gsal
PlanCreationDate | Wednesday, October 02, 2019 09
PlanModificationDate | Wednesday, October 02, 2019 10
StructureId | AirGap3
StructureName | AirGap3
PhotonAlg | AAA_10028_Golden_Beam
CalculationGridSizeInCM | 0.25
FieldNormalizationType | 100
HeterogeneityCorrection | ON
StdArcCalculationSegments | 10
VMATFluenceResolution | High
TargetVolume | -
PlanNormValue | 100
PlanNormMethod | Plan Normalization Value
PrescribedDosePercent | 100
PrimaryRefPointId | -
RelativeDoseInPrimaryRefPoint | -
Fractionations | 1
FractionationId | F1
PrescribedDose | 100
PrescribedDosePerFraction | 100
DoseInPrimaryRefPoint | -
DosePerFractionInPrimaryRefPoint | -
Fractions | 1

## points (2)
Dose | Depth | SSD | Point | Field | Side | Field Side | Field SSD | Point | X | Y | Z
48.1 | 0 | 100 | Thin d=0, 5 | 20x20 6MV 0Nd | Thin | Thin | 100 | Thick d=12, -5 | -7.5 | 9 | 0
81.3 | 3 | 97.5 | Thick d=3, -5 | 20x20 6MV 0Nd | Thick | Thin | 100 | Thick d=3, -5 | -7.5 | 0 | 0 | Point Dose |  |  | Side
154.4 | 3 | 100 | Thin d=3, 5 | 20x20 6MV 0Nd | Thin | Thin | 100 | Thick d=6, -5 | -7.5 | 3 | 0 | Field | Field SSD | Depth | Thick | Thin | Difference
124.2 | 6 | 97.5 | Thick d=6, -5 | 20x20 6MV 0Nd | Thick | Thin | 100 | Thick d=9, -5 | -7.5 | 6 | 0 | 20x20 6MV 0Nd | 100 | 3 | 81.3 | 154.4 | 73.1
136.1 | 6 | 100 | Thin d=6, 5 | 20x20 6MV 0Nd | Thin | Thin | 100 | Thin d=0, 5 | 2.5 | 0 | 0 | 20x20 6MV 0Nd | 100 | 6 | 124.2 | 136.1 | 11.9
112.7 | 9 | 97.5 | Thick d=9, -5 | 20x20 6MV 0Nd | Thick | Thin | 100 | Thin d=3, 5 | 2.5 | 3 | 0 | 20x20 6MV 0Nd | 100 | 9 | 112.7 | 118.4 | 5.7
118.4 | 9 | 100 | Thin d=9, 5 | 20x20 6MV 0Nd | Thin | Thin | 100 | Thin d=6, 5 | 2.5 | 6 | 0 | 20x20 6MV 3Nd | 97 | 3 | 15.9 | 101.2 | 85.3
97.8 | 12 | 97.4 | Thick d=12, -5 | 20x20 6MV 0Nd | Thick | Thin | 100 | Thin d=9, 5 | 2.5 | 9 | 0 | 20x20 6MV 3Nd | 97 | 6 | 46.7 | 88.9 | 42.2
32.2 | 0 | 100 | Thin d=0, 5 | 20x20 6MV 3Nd | Thin | Thin | 97 |  |  |  |  | 20x20 6MV 3Nd | 97 | 9 | 68.5 | 77.2 | 8.7
15.9 | 3 | 94.5 | Thick d=3, -5 | 20x20 6MV 3Nd | Thick | Thin | 97 |  |  |  |  | 20x20 6MV 6Nd | 94 | 3 | 10.6 | 115.2 | 104.6
101.2 | 3 | 97 | Thin d=3, 5 | 20x20 6MV 3Nd | Thin | Thin | 97 |  |  |  |  | 20x20 6MV 6Nd | 94 | 6 | 19 | 101 | 82
46.7 | 6 | 94.4 | Thick d=6, -5 | 20x20 6MV 3Nd | Thick | Thin | 97 |  |  |  |  | 20x20 6MV 6Nd | 94 | 9 | 46.1 | 87.5 | 41.4
88.9 | 6 | 97 | Thin d=6, 5 | 20x20 6MV 3Nd | Thin | Thin | 97 |  |  |  |  | 20x20 6MV 9Nd | 91 | 3 | 9.1 | 133.2 | 124.1
68.5 | 9 | 94.4 | Thick d=9, -5 | 20x20 6MV 3Nd | Thick | Thin | 97 |  |  |  |  | 20x20 6MV 9Nd | 91 | 6 | 14.4 | 116.5 | 102.1
77.2 | 9 | 97 | Thin d=9, 5 | 20x20 6MV 3Nd | Thin | Thin | 97 |  |  |  |  | 20x20 6MV 9Nd | 91 | 9 | 21.3 | 100.6 | 79.3
62.1 | 12 | 94.4 | Thick d=12, -5 | 20x20 6MV 3Nd | Thick | Thin | 97
37.6 | 0 | 100 | Thin d=0, 5 | 20x20 6MV 6Nd | Thin | Thin | 94
10.6 | 3 | 91.5 | Thick d=3, -5 | 20x20 6MV 6Nd | Thick | Thin | 94
115.2 | 3 | 94 | Thin d=3, 5 | 20x20 6MV 6Nd | Thin | Thin | 94
19 | 6 | 91.5 | Thick d=6, -5 | 20x20 6MV 6Nd | Thick | Thin | 94
101 | 6 | 94 | Thin d=6, 5 | 20x20 6MV 6Nd | Thin | Thin | 94
46.1 | 9 | 91.4 | Thick d=9, -5 | 20x20 6MV 6Nd | Thick | Thin | 94
87.5 | 9 | 94 | Thin d=9, 5 | 20x20 6MV 6Nd | Thin | Thin | 94
65.4 | 12 | 91.4 | Thick d=12, -5 | 20x20 6MV 6Nd | Thick | Thin | 94
44.1 | 0 | 100 | Thin d=0, 5 | 20x20 6MV 9Nd | Thin | Thin | 91
9.1 | 3 | 88.5 | Thick d=3, -5 | 20x20 6MV 9Nd | Thick | Thin | 91
133.2 | 3 | 91 | Thin d=3, 5 | 20x20 6MV 9Nd | Thin | Thin | 91
14.4 | 6 | 88.5 | Thick d=6, -5 | 20x20 6MV 9Nd | Thick | Thin | 91
116.5 | 6 | 91 | Thin d=6, 5 | 20x20 6MV 9Nd | Thin | Thin | 91
21.3 | 9 | 88.5 | Thick d=9, -5 | 20x20 6MV 9Nd | Thick | Thin | 91
100.6 | 9 | 91 | Thin d=9, 5 | 20x20 6MV 9Nd | Thin | Thin | 91
45.8 | 12 | 88.4 | Thick d=12, -5 | 20x20 6MV 9Nd | Thick | Thin | 91
271.2 | 0 | 100 | Thick d=12, -5 | Total | Thick | #N/A | #N/A
116.9 | 0 | 100 | Thick d=3, -5 | Total | Thick | #N/A | #N/A
204.3 | 0 | 100 | Thick d=6, -5 | Total | Thick | #N/A | #N/A
248.6 | 0 | 100 | Thick d=9, -5 | Total | Thick | #N/A | #N/A
162 | 0 | 100 | Thin d=0, 5 | Total | Thin | #N/A | #N/A
504 | 0 | 100 | Thin d=3, 5 | Total | Thin | #N/A | #N/A
442.5 | 0 | 100 | Thin d=6, 5 | Total | Thin | #N/A | #N/A
383.7 | 0 | 100 | Thin d=9, 5 | Total | Thin | #N/A | #N/A

## fields (2)
Field | Field Side | Field SSD | Course | Plan | Machine | Energy | Gantry | Collimator | Couch | X1 | X2 | Y1 | Y2 | Iso X | Iso Y | Iso Z | SAD | Mus | Field Weight | Field Dose | Norm Factor | Technique | Warning
20x20 6MV 0Nd | Thin | 100 | AirGap2 | AG30N6XF20 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 0 | 0 | 100 | 157 | 1.634 | 161.7 | 100 | STATIC | Dose in isocenter is too small
20x20 6MV 3Nd | Thin | 97 | AirGap2 | AG30N6XF20 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 3.02 | 0 | 100 | 97 | 1.011 | 106.1 | 100 | STATIC
20x20 6MV 6Nd | Thin | 94 | AirGap2 | AG30N6XF20 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 6.02 | 0 | 100 | 104 | 1.009 | 120.9 | 100 | STATIC
20x20 6MV 9Nd | Thin | 91 | AirGap2 | AG30N6XF20 | TR1 | 6 | 0 | 0 | 0 | 10 | 10 | 10 | 10 | 2.5 | 9.02 | 0 | 100 | 113 | 1.005 | 140.1 | 100 | STATIC