## Item
Item_ID | Item_Name | Item_View | Item_Category | Item_Description | Item_Criteria | Item_Values | Item_Reference_View | Item_Check_Order | Default_Include
168 | Accuracy | Electron & Photon Calculation Models | Calculation Parameters |  | Appropriate | 2% |  | 61 | FALSE
70 | Grid Size | Electron & Photon Calculation Models | Calculation Parameters |  | Appropriate | 0.25 |  | 62 | TRUE
169 | Smoothing | Electron & Photon Calculation Models | Calculation Parameters |  | Appropriate | 3D Gaussian Medium |  | 63 | FALSE
71 | Field Normalization Type | Electron & Photon Calculation Models | Calculation Parameters |  | Appropriate | None |  | 64 | TRUE
72 | Heterogenety | Electron & Photon Calculation Models | Calculation Parameters |  | Appropriate | ON |  | 65 | TRUE
73 | Calculation Model | Electron & Photon Calculation Models | Calculation Parameters |  | Appropriate | AAA_8117_21A_Measured |  | 66 | TRUE
248 | DVO Inhomogeneity | Optimization Calculation Models | Calculation Parameters |  | Appropriate | ON |  | 66.1 | FALSE
249 | DVO SmoothX | Optimize | Calculation Parameters |  | Appropriate | 80 |  | 66.2 | FALSE
250 | DVO SmoothY | Optimize | Calculation Parameters |  | Appropriate | 60 |  | 66.3 | FALSE
251 | DVO Minimize dose | Optimize | Calculation Parameters |  | Appropriate | 0 |  | 66.4 | FALSE
269 | Structure Resolution | Optimize | Calculation Parameters |  | Appropriate |  |  | 66.5 | FALSE
258 | Optimization Parameters | Optimize | Calculation Parameters |  | Appropriate | Constraints are set appropriately |  | 107.5 | FALSE
259 | Optimization Minimization | Optimize | Calculation Parameters |  | Appropriate | The optimization parameter has been appropriately minimized |  | 107.6 | FALSE
187 | Pencil Beam & Grid Size | TPP Dose Distribution | Calculation Parameters |  | Appropriate | 0.6 cm |  | 104 | FALSE
225 | Inhomogeneity | TPP Plan Data Sheet | Calculation Parameters |  | Appropriate | 3D |  | 99 | FALSE
186 | Unit Modeling | TPP Plan Data Sheet | Calculation Parameters |  | Appropriate |  |  | 100 | FALSE
68 | Normalization Mode | Dose Prescription | Calculation Parameters |  | For Information |  |  | 60 | TRUE
159 | Notes | Clinical Progress Notes | Comments |  | For Information |  |  | 8 | TRUE
48 | Notes | Technical Progress Notes | Comments |  | For Information |  |  | 9 | TRUE
227 | Image Registration | CT Volume Orthogonal Views | Contours |  | Appropriate | Matches in GTV region |  | 43 | TRUE
50 | Scanned Region | CT Volume Orthogonal Views | Contours |  | Appropriate | Covers all regions of interest |  | 44 | TRUE
57 | GTV, CTV, PTV | CT Volume Orthogonal Views | Contours |  | Appropriate | Contours and Margins are reasonable |  | 51 | TRUE
58 | Organs at Risk | CT Volume Orthogonal Views | Contours |  | Appropriate | Appropriate organs are contoured |  | 53 | TRUE
52 | Closed Body Contour | CT Volume Orthogonal Views | Contours |  | Correct |  |  | 45 | TRUE
167 | BBs & Wires Contoured out | CT Volume Orthogonal Views | Contours |  | Correct |  |  | 52 | FALSE
156 | Plan Approval date | Final Plan Report | Document Validated |  | Consistent |  | Final Plan Report | 115.7 | TRUE
38 | Date | Preliminary CT Plan Report | Document Validated |  | Consistent |  | CT Simulation Worksheet | 21 | FALSE
63 | Status | CT Volume Properties - User | Document Validated |  | Correct |  |  | 42 | TRUE
98 | Plan Approval | Plan Properties - General | Document Validated |  | Correct |  |  | 112 | TRUE
99 | Exported to IMSure | Plan Properties - General | Document Validated |  | Correct |  |  | 113 | TRUE
184 | Signed by Dosimetrist | TPP Plan Data Sheet | Document Validated |  | Correct |  |  | 97 | FALSE
266 | Task Completed | Activity Capture | Document Validated |  | Entered |  |  | 176 | TRUE
256 | Workload Entered | Activity Capture | Document Validated |  | Entered |  |  | 177 | TRUE
257 | Chart Tracked | CRTP | Document Validated |  | Entered |  |  | 178 | TRUE
274 | IMRT EPID Record | IMRT EPID Spreadsheet | Document Validated |  | Entered |  |  | 179 | FALSE
158 | Dynamic Document | IMSure Photons & Electrons | Document Validated |  | Entered |  |  | 137 | TRUE
217 | Calculation Printed and Signed | MU Calc | Document Validated |  | Entered |  |  | 146 | FALSE
194 | Complete QA Checklist | QA Task | Document Validated |  | Entered |  |  | 174 | TRUE
195 | Complete & Sign QA checklist | QA task sheet | Document Validated |  | Entered |  |  | 175 | FALSE
150 | Treatment Approval | RT Chart Reference Points | Document Validated |  | Entered |  |  | 173 | TRUE
9 | Signed | Prescription Sheet | Document Validated |  | Present |  |  | 7 | TRUE
120 | Hot Spots | Dose Distributions | Dose data |  | Appropriate | Uniform Distribution |  | 84 | TRUE
95 | Isodose Levels | Dose Distributions | Dose data |  | Appropriate |  |  | 85 | FALSE
93 | Calculation Volume | DVH | Dose data |  | Appropriate | Includes All Relevant Structures |  | 108 | TRUE
92 | PTV and CTV coverage | DVH | Dose data |  | Appropriate | PTV gets 90% CTV gets 95% |  | 109 | TRUE
232 | Gross Disease (PTV70) | DVH | Dose data |  | Appropriate | D99% ≥ 65.1 Gy, D95% ≥ 70 Gy V77Gy ≤ 2% V80.5Gy ≤0% |  | 109.1 | FALSE
260 | Gross Disease (PTV76) | DVH | Dose data |  | Appropriate | D98% ≥ 76 Gy V81Gy ≤ 2% V83Gy ≤0% |  | 109.11 | FALSE
261 | (CTV76) | DVH | Dose data |  | Appropriate | D100% ≥ 76 Gy |  | 109.12 | FALSE
262 | (CTV70) | DVH | Dose data |  | Appropriate | D100% ≥ 70 Gy |  | 109.13 | FALSE
235 | Residual Disease (PTV66) | DVH | Dose data |  | Appropriate | D99% ≥ 61.4 Gy, D95% ≥ 66 Gy V72.6Gy ≤ 2% V75.9Gy ≤0% |  | 109.15 | FALSE
233 | Intermediate Risk (PTV63) | DVH | Dose data |  | Appropriate | D99% ≥ 58.6 Gy, D95% ≥ 63 Gy V77Gy ≤ 0% |  | 109.2 | FALSE
234 | Intermediate Risk (PTV60) | DVH | Dose data |  | Appropriate | D99% ≥ 55.8 Gy, D95% ≥ 60 Gy V72.6Gy ≤ 0%-Resiual Disease V66Gy ≤ 2%,V69Gy ≤ 0%-No Resiual Disease |  | 109.3 | FALSE
267 | Elective Nodes (PTV56) | DVH | Dose data |  | Appropriate | D99% ≥ 53.2 Gy, D95% ≥ 56 Gy V77Gy ≤ 0% |  | 109.35 | FALSE
237 | Elective Nodes (PTV54) | DVH | Dose data |  | Appropriate | D99% ≥ 53.2 Gy, D95% ≥ 54 Gy V72.6Gy ≤ 0% |  | 109.4 | FALSE
96 | Dose to Critical Organs | DVH | Dose data |  | Appropriate | Based on Technique |  | 110 | TRUE
238 | Parotid & Submandibulars | DVH | Dose data |  | Appropriate | Rt or Lt V30Gy ≤ 50% Rt or Lt Mean Dose ≤ 26Gy Combined Rt & Lt V20Gy ≤ 20cc |  | 110.1 | FALSE
239 | Mandible | DVH | Dose data |  | Appropriate | V70Gy ≤ 0.1cc |  | 110.2 | FALSE
240 | Cord | DVH | Dose data |  | Appropriate | V45Gy ≤ 0.1cc, V48Gy ≤ 0cc |  | 110.3 | FALSE
263 | Bladder | DVH | Dose data |  | Appropriate | V80Gy ≤ 15%, V75Gy ≤ 25%, V70Gy ≤ 35%, V50Gy ≤ 65% |  | 110.31 | FALSE
264 | Rectum | DVH | Dose data |  | Appropriate | V75Gy ≤ 15%, V70Gy ≤ 25%, V65Gy ≤ 35%, V60Gy ≤ 50% |  | 110.32 | FALSE
265 | Femural Head | DVH | Dose data |  | Appropriate | V52Gy ≤ 0% |  | 110.33 | FALSE
241 | Brain Stem | DVH | Dose data |  | Appropriate | V50Gy ≤ 0.1cc, V54Gy ≤ 0cc |  | 110.4 | FALSE
242 | Chiasm | DVH | Dose data |  | Appropriate | V50Gy ≤ 0% |  | 110.5 | FALSE
243 | Lt & Rt Optic Nerve | DVH | Dose data |  | Appropriate | V50Gy ≤ 0% |  | 110.6 | FALSE
244 | Lt & Rt Orbit | DVH | Dose data |  | Appropriate | V45Gy ≤ 0% |  | 110.7 | FALSE
272 | Lens | DVH | Dose data |  | Appropriate | V10Gy ≤ 0% |  | 110.75 | FALSE
245 | Larynx | DVH | Dose data |  | Appropriate | V50Gy ≤ 67% Mean Dose ≤ 45Gy |  | 110.8 | FALSE
268 | Brachial Plexus | DVH | Dose data |  | Appropriate |  |  | 110.9 | FALSE
253 | Reconstructed PI dose | EPID Dose QA Report | Dose data |  | Appropriate | > 3% difference |  | 173.1 | FALSE
255 | Histogram | EPID Dose QA Report | Dose data |  | Appropriate | 85% ¦CHI¦ <1 |  | 173.2 | FALSE
100 | Equivalent Path Length | Final Plan Report | Dose data |  | Appropriate | Reasonable values |  | 116 | TRUE
97 | MUREF Location | Reference Point Properties - Location | Dose data |  | Appropriate | Near Center of field, Away from Inhomogeneities |  | 111 | TRUE
189 | Dose coverage | TPP Dose Distribution | Dose data |  | Appropriate | 93-109% (Prefer 95-107%) |  | 107 | FALSE
190 | Normalization | TPP Plan Data Sheet | Dose data |  | Appropriate | Seems Reasonable |  | 98 | FALSE
118 | MU Calculation | IMSure Photons & Electrons | Dose data |  | Correct |  |  | 136 | TRUE
254 | IMRT Matrixx QA Results | Matrixx Dose QA Report | Dose data |  | Correct |  |  | 173.3 | FALSE
216 | MU Calculation | MU Calc | Dose data |  | Correct |  |  | 145 | FALSE
110 | Effective Depth | IMSure Calc Points | Dose data |  | Entered |  | Final Plan Report | 125 | TRUE
117 | User factor | IMSure Photons & Electrons | Dose data |  | Entered |  |  | 135 | TRUE
163 | MU_on & MU_off | IMSure User Factor | Dose data |  | Entered |  |  | 134 | TRUE
203 | Dose & Fractionation | MU Calc | Dose data |  | Entered |  |  | 142 | FALSE
205 | Normalization Factor | MU Calc | Dose data |  | Entered |  |  | 144 | FALSE
90 | MLC Coverage | BEVs | Field Parameters |  | Appropriate | Margin around PTV is reasonable |  | 87 | TRUE
252 | Fluences | BEVs | Field Parameters |  | Appropriate | Fluence Gradients Look Reasonable |  | 87.5 | FALSE
153 | Breast Skin Flash | BEVs | Field Parameters |  | Appropriate | 2 cm |  | 88 | FALSE
152 | Field-in-Field | BEVs | Field Parameters |  | Appropriate | Subfields big to small, Smallest subfield > 1 cm |  | 89 | FALSE
91 | MLC Leaf Gap | BEVs | Field Parameters |  | Appropriate | Minimized by Jaw settings, Leafs Not Closed in Field |  | 90 | TRUE
276 | Field-in-Field MUs | BEVs | Field Parameters |  | Appropriate | MUs > 5 per sub field |  | 90.5 | FALSE
171 | Field-in-Field Matching | BEVs | Field Parameters |  | Appropriate | Matches with Individual Plan Field Shapes |  | 91 | FALSE
178 | Block Tray | BEVs in TPP | Field Parameters |  | Appropriate | None - 40 cm |  | 101 | FALSE
179 | Block Transmission | BEVs in TPP | Field Parameters |  | Appropriate | 3.0-6MV, 2.0-4MV |  | 102 | FALSE
77 | Beam Energy | Field Parameters | Field Parameters |  | Appropriate | Based on Technique |  | 73 | TRUE
78 | Beam Weight | Field Parameters | Field Parameters |  | Appropriate | Adds up to 1 |  | 74 | FALSE
79 | Gantry and Couch Rotation | Field Parameters | Field Parameters |  | Appropriate | Based on Technique |  | 75 | TRUE
80 | Collimator Rotation | Field Parameters | Field Parameters |  | Appropriate | For Field Matching and Dynamic Wedge |  | 76 | TRUE
82 | Wedges | Field Parameters | Field Parameters |  | Appropriate | Based on Technique |  | 77 | TRUE
94 | Field Jaw Settings | Field Parameters | Field Parameters |  | Appropriate |  |  | 78 | TRUE
83 | Field Size | Field Parameters | Field Parameters |  | Appropriate | Jaw Values are Rounded Correctly |  | 79 | FALSE
86 | MUs | Field Parameters | Field Parameters |  | Appropriate | Values are reasonable |  | 83 | TRUE
129 | Technique | RT Chart Field Parameters | Field Parameters |  | Appropriate | Static |  | 153 | TRUE
131 | Dose Rate, Time | RT Chart Field Parameters | Field Parameters |  | Appropriate | 400, 1.00 min |  | 154 | TRUE
134 | Tolerance Table | RT Chart Field Parameters | Field Parameters |  | Appropriate | Photon |  | 155 | TRUE
157 | 180 E | RT Chart Field Parameters | Field Parameters |  | Appropriate | 180 E where appropriate |  | 156.5 | TRUE
139 | MLC | RT Chart Field Parameters | Field Parameters |  | Appropriate | Static | Field Parameters | 160 | TRUE
142 | Couch Lng & Imager Vrt | RT Chart Field Parameters | Field Parameters |  | Appropriate | 100,50 |  | 165 | TRUE
144 | Setup Notes | RT Chart Field Parameters | Field Parameters |  | Appropriate | Indicate Bolus |  | 167 | TRUE
188 | Wedges | TPP Dose Distribution | Field Parameters |  | Appropriate |  |  | 105 | FALSE
181 | Beam Weighting | TPP Plan Data Sheet | Field Parameters |  | Appropriate | Adds up to 1.00 |  | 94 | FALSE
154 | Field Parameters | Dose Distribution | Field Parameters |  | Consistent |  | Field Parameters | 114 | TRUE
108 | CAX SSD | IMSure Calc Points | Field Parameters |  | Consistent |  | Final Plan Report | 122 | FALSE
161 | Dose Per Field | IMSure Calc Points | Field Parameters |  | Consistent |  | Final Plan Report | 123 | FALSE
105 | Treatment Unit | IMSure Photons & Electrons | Field Parameters |  | Consistent |  | Field Parameters | 126 | FALSE
111 | Energy | IMSure Photons & Electrons | Field Parameters |  | Consistent |  | Field Parameters | 128 | FALSE
112 | Wedges | IMSure Photons & Electrons | Field Parameters |  | Consistent |  | Field Parameters | 129 | TRUE
113 | Angles and Field Size | IMSure Photons & Electrons | Field Parameters |  | Consistent |  | Field Parameters | 130 | FALSE
164 | Eclipse MUs | IMSure Photons & Electrons | Field Parameters |  | Consistent |  | Field Parameters | 131 | TRUE
160 | Field Shapes | IMSure Photons & Electrons | Field Parameters |  | Consistent |  | BEVs | 132 | FALSE
166 | MLC Field Shape | Initial Multi-Field Plan | Field Parameters |  | Consistent |  | BEVs | 92 | FALSE
193 | Portal Image Isocentre | Portal Image Plans | Field Parameters |  | Consistent |  | Field Parameters | 148 | FALSE
127 | Treatment Machine | RT Chart Field Parameters | Field Parameters |  | Consistent |  | Field Parameters | 151 | TRUE
128 | Field ID | RT Chart Field Parameters | Field Parameters |  | Consistent |  | Field Parameters | 152 | TRUE
220 | Beam Energy | RT Chart Field Parameters | Field Parameters |  | Consistent |  | TPP Dose Distribution | 153.5 | FALSE
130 | MU | RT Chart Field Parameters | Field Parameters |  | Consistent |  | Field Parameters | 154.5 | TRUE
228 | SSD | RT Chart Field Parameters | Field Parameters |  | Consistent |  | Field Parameters | 155.5 | TRUE
219 | Gantry, Collimator, Couch angles | RT Chart Field Parameters | Field Parameters |  | Consistent |  | TPP Dose Distribution | 156 | FALSE
226 | Field Size | RT Chart Field Parameters | Field Parameters |  | Consistent |  | TPP Dose Distribution | 157 | FALSE
229 | Wedges | RT Chart Field Parameters | Field Parameters |  | Consistent |  | Field Parameters | 161 | TRUE
221 | Block Tray | RT Chart Field Parameters | Field Parameters |  | Consistent |  | TPP Dose Distribution | 163 | FALSE
141 | Bolus | RT Chart Field Parameters | Field Parameters |  | Consistent |  | Dose Distribution | 164 | TRUE
180 | Beam Energy | TPP Plan Data Sheet | Field Parameters |  | Consistent |  | Field Parameters | 93 | FALSE
182 | Angles, Field Size, SSD, Wedges, Bolus | TPP Plan Data Sheet | Field Parameters |  | Consistent |  | Field Parameters | 95 | FALSE
223 | Isocentre Location | TPP Plan Data Sheet | Field Parameters |  | Consistent |  | Eclipse and TPP Isocentre QA | 96 | FALSE
75 | Field Names | Field Parameters | Field Parameters |  | Correct |  |  | 71 | TRUE
76 | Treatment Units The Same | Field Parameters | Field Parameters |  | Correct |  |  | 72 | TRUE
165 | Validate Plan | RT Chart Field Parameters | Field Parameters |  | Correct |  |  | 151.3 | TRUE
109 | PSSD | IMSure Calc Points | Field Parameters |  | Entered |  | Final Plan Report | 124 | TRUE
199 | Plan Datasheet copied & Pasted | MU Calc | Field Parameters |  | Entered |  |  | 140 | FALSE
202 | Bolus Thickness | MU Calc | Field Parameters |  | Entered |  |  | 141 | FALSE
204 | Block/MLC | MU Calc | Field Parameters |  | Entered |  |  | 143 | FALSE
145 | Graticle & Apperature Contours | RT Chart Field Parameters | Field Parameters |  | Present |  |  | 151.5 | TRUE
7 | Patient ID sticker | Technical Progress Notes | Patient ID |  | Appropriate | Agrees with Patient Manager |  | 8.9 | TRUE
10 | Patient Name and ID | CT Simulation Worksheet | Patient ID |  | Consistent |  | Prescription Sheet | 10 | TRUE
172 | Patient Name & ID | Field Matching Worksheet | Patient ID |  | Consistent |  | CT Simulation Worksheet | 54 | FALSE
101 | Patient Name and ID | IMSure Calc Points | Patient ID |  | Consistent |  | Prescription Sheet | 117 | TRUE
33 | Patient Name and ID | LapData Report | Patient ID |  | Consistent |  | CT Simulation Worksheet | 29 | TRUE
22 | Patient Name & ID | Preliminary CT Plan Report | Patient ID |  | Consistent |  | CT Simulation Worksheet | 22 | TRUE
6 | Patient Name | Prescription Sheet | Patient ID |  | Correct |  |  | 2 | TRUE
197 | Patient Name and ID | MU Calc | Patient ID |  | Entered |  |  | 138 | FALSE
122 | Plan Label | IMSure Calc Points | Plan ID |  | Consistent |  | Final Plan Report | 118 | TRUE
121 | Plan Name | RT Chart Field Parameters | Plan ID |  | Consistent |  | CT Volume Orthogonal Views | 150 | TRUE
185 | Plan Name & Date | TPP Dose Distribution | Plan ID |  | Consistent |  | TPP Plan Data Sheet | 103 | FALSE
162 | Patient & Beam Info | IMSure User Factor | Plan ID |  | Entered |  |  | 133 | TRUE
198 | Plan Label | MU Calc | Plan ID |  | Entered |  |  | 139 | FALSE
1 | Site | Course | Prescription |  | Appropriate | Consistant with Diagnosis |  | 4 | TRUE
2 | Dose and Fractions | Dose Prescription | Prescription |  | Appropriate |  |  | 6 | TRUE
103 | Isodose Line | IMSure Calc Points | Prescription |  | Appropriate | 100 |  | 119 | FALSE
230 | Bolus | Prescription Sheet | Prescription |  | Appropriate |  |  | 5 | TRUE
148 | Primary Reference Point | RT Chart Reference Points | Prescription |  | Appropriate | Dose limits add up correctly |  | 171 | TRUE
149 | MUREF | RT Chart Reference Points | Prescription |  | Appropriate | Dose limit ≤ primary reference |  | 172 | TRUE
66 | Dose and Fractions | Dose Prescription | Prescription |  | Consistent |  | Prescription Sheet | 59 | FALSE
123 | Dose & Fractionation | RT Chart Prescription | Prescription |  | Consistent |  | Dose Prescription | 149 | TRUE
270 | Scheduled Fractions | Scheduling | Prescription |  | Consistent |  | RT Chart Prescription | 168 | TRUE
271 | Diagnosis | Course | Prescription |  | For Information |  |  | 3.5 | TRUE
89 | BEV Image Quality | BEVs | Scan Data |  | Appropriate | Landmarks are Visable |  | 86 | TRUE
44 | Slice Thickness | CT Image Properites - CT Geometry | Scan Data |  | Appropriate |  |  | 37 | TRUE
39 | Contrast | CT Simulation Worksheet | Scan Data |  | Appropriate |  |  | 16 | TRUE
62 | Plane Separation | CT Volume Properties - 3D | Scan Data |  | Appropriate |  |  | 38 | TRUE
42 | Series ID | CT Image Properites - Technical | Scan Data |  | Consistent |  | Preliminary CT Plan Report | 34 | TRUE
59 | Number of CT Slices | CT Volume Properties - 3D | Scan Data |  | Consistent |  | CT Simulation Worksheet | 39 | TRUE
46 | Volume ID | CT Volume Properties - User | Scan Data | OPIS Code for treatment Site | Consistent |  | CT Simulation Worksheet | 41 | TRUE
35 | Run Number | LapData Report | Scan Data |  | Consistent |  | CT Simulation Worksheet | 31 | TRUE
41 | Image Series ID | Preliminary CT Plan Report | Scan Data | contains CT run number and Site code | Consistent |  | CT Simulation Worksheet | 23 | FALSE
40 | Device | CT Image Properites - Technical | Scan Data |  | Correct |  |  | 35 | TRUE
13 | CT Scan Study # | CT Simulation Worksheet | Scan Data |  | Present |  |  | 11 | TRUE
17 | Patient Orientation | CT Simulation Worksheet | Set-up |  | Appropriate | Head First Supine |  | 14 | TRUE
19 | Immobilization | CT Simulation Worksheet | Set-up |  | Appropriate |  |  | 17 | TRUE
20 | Lat. Tattoo Reference | CT Volume Orthogonal Views | Set-up |  | Appropriate | At Level of AP Straightening tattoo |  | 46.1 | TRUE
222 | Straightening tattoo | CT Volume Orthogonal Views | Set-up |  | Appropriate | 10 cm SUP / INF of AP tattoo |  | 46.2 | TRUE
56 | Contrast | CT Volume Orthogonal Views | Set-up |  | Appropriate |  |  | 49 | TRUE
146 | Patient Anatomy | CT Volume Orthogonal Views | Set-up |  | Appropriate | No CT or physical artifacts |  | 50 | TRUE
170 | Applicator Clearance | Field Parameters | Set-up |  | Appropriate | Min 95 cm clearance |  | 80 | FALSE
84 | Isocentre Shifts | Field Parameters | Set-up |  | Appropriate |  |  | 81 | TRUE
85 | SSD | Field Parameters | Set-up |  | Appropriate | Reasonable, Centred POPs |  | 82 | TRUE
27 | Treatment Orientation | Preliminary CT Plan Report | Set-up |  | Appropriate | HFS |  | 24.5 | TRUE
28 | Image Origin | Preliminary CT Plan Report | Set-up |  | Appropriate | X,Z small Y -25 to -35 |  | 25 | TRUE
43 | Table Height | CT Image Properites - CT Geometry | Set-up |  | Consistent |  | CT Simulation Worksheet | 36 | TRUE
16 | Treatment Sites | CT Simulation Worksheet | Set-up | OPIS Code for treatment Site | Consistent |  | Prescription Sheet | 12 | TRUE
65 | TTH Measured | CT Volume Orthogonal Views | Set-up |  | Consistent |  | CT Simulation Worksheet | 46.5 | FALSE
45 | Offset from DICOM origin | CT Volume Properties - Origin | Set-up |  | Consistent |  | Preliminary CT Plan Report | 40 | TRUE
155 | ICHAT & Shifts | Dose Distribution | Set-up |  | Consistent |  | Axial Slices | 115 | TRUE
231 | SSD | Final Plan Report | Set-up |  | Consistent |  | Field Parameters | 115.5 | FALSE
107 | MUREF Location | IMSure Calc Points | Set-up |  | Consistent |  | Eclipse and TPP Isocentre QA | 121 | TRUE
106 | Isocentre | IMSure Photons & Electrons | Set-up |  | Consistent |  | Eclipse and TPP Isocentre QA | 127 | FALSE
34 | Isocentre | LapData Report | Set-up |  | Consistent |  | Eclipse and TPP Isocentre QA | 30 | TRUE
36 | CT SIM Iso | LapData Report | Set-up |  | Consistent |  | CT Simulation Worksheet | 32 | TRUE
61 | Treatment Site | LapData Report | Set-up |  | Consistent |  | CT Simulation Worksheet | 33 | TRUE
26 | Imaging Orientation | Preliminary CT Plan Report | Set-up |  | Consistent |  | CT Simulation Worksheet | 24 | FALSE
224 | Set-up Reference Labels | TPP Dose Distribution | Set-up |  | Consistent |  | Axial Slices | 106 | FALSE
53 | BBs on Origin | CT Volume Orthogonal Views | Set-up |  | Correct |  |  | 46 | TRUE
51 | Patient Alignment | CT Volume Orthogonal Views | Set-up |  | Correct |  |  | 47 | FALSE
29 | Original Couch Vertical | Eclipse and TPP Isocentre QA | Set-up |  | Entered |  | CT Simulation Worksheet | 27 | TRUE
30 | Image Origin | Eclipse and TPP Isocentre QA | Set-up |  | Entered |  | Preliminary CT Plan Report | 28 | TRUE
273 | Isocentre | Eclipse and TPP Isocentre QA | Set-up |  | Entered |  | Field Parameters | 28.5 | TRUE
104 | Patient Origin | IMSure Calc Points | Set-up |  | Entered | Eclipse and TPP Isocentre QA | Eclipse and TPP Isocentre QA | 120 | TRUE
15 | Original CT Couch Height | CT Simulation Worksheet | Set-up |  | Present |  |  | 13 | TRUE
18 | Number of CT Slices | CT Simulation Worksheet | Set-up |  | Present |  |  | 15 | TRUE
21 | TTH | CT Simulation Worksheet | Set-up |  | Present |  |  | 20 | TRUE
88 | Bolus Thickness and Extent | Axial Slices | Treatment Technique |  | Appropriate | Based on Technique |  | 68 | TRUE
64 | intent | Course | Treatment Technique |  | Appropriate | Curative / Adjuvent / Palliative |  | 3 | TRUE
175 | Initial Collimator Angle | Field Matching Worksheet | Treatment Technique |  | Appropriate | Based on patient anatomy |  | 57 | FALSE
176 | Final Collimator & Couch angles | Field Matching Worksheet | Treatment Technique |  | Appropriate | Seem Reasonable |  | 58 | FALSE
74 | Beam Arrangement | Field Parameters | Treatment Technique |  | Appropriate |  |  | 67 | TRUE
87 | Bolus Linked to Field | Field Parameters | Treatment Technique |  | Appropriate | Linked to the Right Fields |  | 70 | TRUE
177 | Shielding Blocks | Field Parameters | Treatment Technique |  | Appropriate | Line up with Match Plane Contour |  | 70.5 | FALSE
173 | Gantry | Field Matching Worksheet | Treatment Technique |  | Consistent |  | Axial Slices | 55 | FALSE
174 | Field Length | Field Matching Worksheet | Treatment Technique |  | Consistent |  | Axial Slices | 56 | FALSE

## Technique
Technique_ID | Technique_Disease | Technique_Item | Technique_show | Technique_Data | Technique_criteria | Technique_reference_view
29491 | Brain | Beam Energy | TRUE | 6X | Appropriate
29487 | Brain | Straightening tattoo | FALSE |  | Appropriate
29488 | Brain | Lat. Tattoo Reference | FALSE | Marks on Mask | Appropriate
29489 | Brain | Beam Arrangement | TRUE | POP or 3 fld | Appropriate
29599 | Brain | MLC | TRUE | Static or Dose Dynamic if Field-in-Field | Appropriate
29598 | Brain | Field-in-Field MUs | TRUE |  | Appropriate
29597 | Brain | MLC Leaf Gap | TRUE |  | Appropriate
29596 | Brain | Field-in-Field Matching | TRUE |  | Appropriate
29485 | Brain | Dose and Fractions | TRUE | 25/10, 30/15 PCI, 20/5 palliative, 60/30 or 40/15 Glioma | Appropriate
29490 | Brain | Isocentre | TRUE | No shifts OR Adjusted to reduce dose to eyes | Appropriate
29484 | Brain | Site | TRUE | BRAI | Appropriate
29492 | Brain | Slice Thickness | TRUE | 5 mm | Appropriate
29493 | Brain | Plane Separation | TRUE | 0.5 cm | Appropriate
29495 | Brain | Contrast | FALSE |  | Appropriate
29496 | Brain | Patient & Beam Info | FALSE |  | Appropriate
29497 | Brain | MU_on & MU_off | FALSE |  | Appropriate
29498 | Brain | User factor | FALSE |  | Appropriate
29499 | Brain | Contrast | FALSE |  | Appropriate
29595 | Brain | Field-in-Field | TRUE |  | Appropriate
29486 | Brain | Immobilization | TRUE | Barker Board | Appropriate
23 | Breast Tangents | Bolus Thickness and Extent | TRUE | 0.5 cm on alternate days | Appropriate
28957 | Breast Tangents | Wedges | FALSE |  | Appropriate
31 | Breast Tangents | Beam Energy | TRUE | 6 MV or 6 & 15 MV | Appropriate
30 | Breast Tangents | GTV, CTV, PTV | FALSE |  | Appropriate
29 | Breast Tangents | Patient Alignment | TRUE | Sup & Inf Sagital BB aligned with Isocentre | Appropriate
28 | Breast Tangents | Lat. Tattoo Reference | FALSE |  | Appropriate
27 | Breast Tangents | Immobilization | TRUE | MedTec Breast Board Correct Side | Appropriate
33 | Breast Tangents | Collimator Rotation | TRUE | Opposed, Minimizes Lung in field | Appropriate
25 | Breast Tangents | Field-in-Field | TRUE | Subfields big to small, Smallest subfield > 1 cm | Appropriate
34 | Breast Tangents | Wedges | FALSE |  | Appropriate
22 | Breast Tangents | Contrast | FALSE |  | Appropriate
302 | Breast Tangents | Straightening tattoo | FALSE | Reference to SSN | Appropriate
6 | Breast Tangents | Patient Anatomy | FALSE |  | Appropriate
5 | Breast Tangents | Plane Separation | TRUE | 0.3 cm | Appropriate
4 | Breast Tangents | Slice Thickness | TRUE | 5 mm Thick | Appropriate
3 | Breast Tangents | Contrast | FALSE |  | Appropriate
26 | Breast Tangents | Breast Skin Flash | TRUE | 2 cm | Appropriate
70 | Breast Tangents | Calculation Volume | FALSE |  | Appropriate
196 | Breast Tangents | Portal Image Isocentre | FALSE |  | Appropriate
28950 | Breast Tangents | MU_on & MU_off | FALSE |  | Appropriate
28949 | Breast Tangents | Patient & Beam Info | FALSE |  | Appropriate
28948 | Breast Tangents | Beam Arrangement | TRUE | Matched Post Borders | Appropriate
28947 | Breast Tangents | Image Registration | FALSE |  | Appropriate
73 | Breast Tangents | User factor | TRUE | 0.88 | Appropriate
32 | Breast Tangents | Gantry and Couch Rotation | TRUE | opposed gantry rotation, Fields clear arm | Appropriate
71 | Breast Tangents | Dose to Critical Organs | FALSE |  | Appropriate
2 | Breast Tangents | Site | TRUE | BRER or BREL | Appropriate
41 | Breast Tangents | 180 E | FALSE |  | Appropriate
40 | Breast Tangents | MLC | TRUE | Dose Dynamic | Appropriate
39 | Breast Tangents | Tolerance Table | TRUE | Tangent Breast | Appropriate
38 | Breast Tangents | Wedges | FALSE |  | Appropriate
37 | Breast Tangents | Isodose Line | FALSE |  | Appropriate
36 | Breast Tangents | MLC Coverage | FALSE |  | Appropriate
35 | Breast Tangents | PTV and CTV coverage | FALSE |  | Appropriate
72 | Breast Tangents | Field-in-Field Matching | TRUE | Matches with Individual Plan Field Shapes | Appropriate
250 | Breast Tangents | Organs at Risk | FALSE |  | Appropriate
29138 | Breast Tangents | Field-in-Field MUs | TRUE | MUs > 5 per sub field | Appropriate
1 | Breast Tangents | Dose and Fractions | TRUE | 40/16 or 50/25 | Appropriate
251 | Breast Tangents | Isocentre Shifts | TRUE | BRER: -X, BREL +X, Y 2 to 10 | Appropriate
249 | Breast Tangents | Image Origin | TRUE | BRE: X & Z small, Y -25 to -45 | Appropriate
29058 | Breast with Nodes | Field-in-Field | TRUE | Subfields big to small, Smallest subfield > 1 cm | Appropriate
29057 | Breast with Nodes | Breast Skin Flash | TRUE | 2 cm | Appropriate
29055 | Breast with Nodes | Wedges | TRUE | SCN only if Y < 10 cm otherwise F-in-F | Appropriate
28955 | Breast with Nodes | Image Registration | FALSE |  | Appropriate
29060 | Breast with Nodes | Isodose Line | FALSE |  | Appropriate
29051 | Breast with Nodes | 180 E | FALSE |  | Appropriate
303 | Breast with Nodes | Isocentre Shifts | TRUE | R:-X, L:+X, Y 2 to 10 SCN: on match plane | Appropriate
243 | Breast with Nodes | Lat. Tattoo Reference | FALSE |  | Appropriate
299 | Breast with Nodes | 218 | TRUE |  | Consistent | TPP Plan Data Sheet
29157 | Breast with Nodes | MLC Coverage | TRUE | SCN Lung Vol < 5 cm | Appropriate
29158 | Breast with Nodes | Field-in-Field MUs | TRUE | MU > 5 per sub field | Appropriate
29053 | Breast with Nodes | Normalization Mode | TRUE | BRE or CHW: Reasonable Value SCN: normalized to prescription point | Appropriate
29065 | Breast with Nodes | Field Jaw Settings | TRUE | BRE: Same fieldsize SCN: Half beam | Appropriate
192 | Breast with Nodes | Closed Body Contour | FALSE |  | Appropriate
193 | Breast with Nodes | Tolerance Table | TRUE | 3&4 Field Breast, Photon for S'clav | Appropriate
29159 | Breast with Nodes | Wedges | FALSE |  | Appropriate
197 | Breast with Nodes | Patient Name and ID | FALSE |  | Entered
198 | Breast with Nodes | Plan Label | FALSE |  | Entered
199 | Breast with Nodes | Plan Datasheet copied & Pasted | FALSE |  | Entered
28952 | Breast with Nodes | Bolus | TRUE | 1/c cm Alternate Day | Appropriate
201 | Breast with Nodes | Dose & Fractionation | FALSE |  | Entered
28954 | Breast with Nodes | Image Origin | TRUE | SCN: | Appropriate
202 | Breast with Nodes | Block/MLC | FALSE |  | Entered
203 | Breast with Nodes | Normalization Factor | FALSE |  | Entered
204 | Breast with Nodes | MU Calculation | FALSE |  | Correct
205 | Breast with Nodes | Calculation Printed and Signed | FALSE |  | Entered
206 | Breast with Nodes | Treatment Machine | FALSE |  | Appropriate
211 | Breast with Nodes | Portal Image Isocentre | FALSE |  | Appropriate
29061 | Breast with Nodes | MLC | TRUE | Dose Dynamic, Static for SCN if Wedges | Appropriate
200 | Breast with Nodes | Bolus Thickness | FALSE |  | Entered
28961 | Breast with Nodes | Image Origin | TRUE | BRE: | Appropriate
297 | Breast with Nodes | Inhomogeneity | FALSE | 3D | Appropriate
296 | Breast with Nodes | Set-up Reference Labels | FALSE |  | Consistent | Axial Slices
295 | Breast with Nodes | Isocentre Location | FALSE |  | Consistent | Eclipse and TPP Isocentre QA
294 | Breast with Nodes | Straightening tattoo | FALSE | Reference to SSN | Appropriate
293 | Breast with Nodes | Dose and Fractions | TRUE | 40/16, 50/25 or 60/30 | Appropriate
226 | Breast with Nodes | Collimator Rotation | TRUE | BRE: Small, SCN:0, or 270 if wedges | Appropriate
29059 | Breast with Nodes | Field-in-Field Matching | TRUE | Matches with Individual Plan Field Shapes | Appropriate
182 | Breast with Nodes | User factor | TRUE | BRE: 0.88, SCN:1.0 | Appropriate
228 | Breast with Nodes | GTV, CTV, PTV | TRUE | Match Plane and Baseline Contoured | Appropriate
236 | Breast with Nodes | Number of CT Slices | TRUE |  | Consistent
218 | Breast with Nodes | Gantry and Couch Rotation | TRUE | BRE:Field borders match with baseline and match plane SCN: Gantry: 15/345 | Appropriate
219 | Breast with Nodes | Beam Energy | TRUE | 6 MV or 6 MV & 15 MV | Appropriate
222 | Breast with Nodes | SSD | TRUE | BRE: ~equal SCN: 98.5 or midplane | Appropriate
223 | Breast with Nodes | Calculation Volume | FALSE |  | Appropriate
224 | Breast with Nodes | PTV and CTV coverage | FALSE |  | Appropriate
28960 | Breast with Nodes | Bolus Thickness and Extent | TRUE | 0.5 cm on alternate days | Consistent | Prescription Sheet
237 | Breast with Nodes | Slice Thickness | TRUE | 5 mm | Appropriate
29062 | Breast with Nodes | BBs on Origin | TRUE | BRE: on Medial BB | Appropriate
29063 | Breast with Nodes | BBs on Origin | TRUE | SCN: on SUP Sgital BB | Appropriate
29064 | Breast with Nodes | MLC Coverage | TRUE | BRE:Matches baseline SCN:cord & humeral head shield | Appropriate
242 | Breast with Nodes | Contrast | FALSE |  | Appropriate
241 | Breast with Nodes | Site | TRUE | BRE or CHW, SCN L or R | Appropriate
240 | Breast with Nodes | Immobilization | TRUE | Medtec Board Correct Side 4-fld: Double MBS board | Appropriate
28962 | Breast with Nodes | Treatment Orientation | FALSE |  | Appropriate
28956 | Breast with Nodes | Hot Spots | TRUE | BRE: Dose to breast between 95% and 105% | Appropriate
227 | Breast with Nodes | Organs at Risk | FALSE |  | Appropriate
235 | Breast with Nodes | Plane Separation | TRUE | 0.3 cm | Appropriate
233 | Breast with Nodes | Offset from DICOM origin | TRUE |  | Consistent | Preliminary CT Plan Report
232 | Breast with Nodes | Offset from DICOM origin | TRUE |  | Consistent | Preliminary CT Plan Report
231 | Breast with Nodes | Contrast | FALSE |  | Appropriate
230 | Breast with Nodes | Patient Alignment | TRUE | Sup & Inf Sagital BB aligned with Isocentre | Appropriate
229 | Breast with Nodes | Patient Anatomy | FALSE |  | Appropriate
225 | Breast with Nodes | Dose to Critical Organs | FALSE |  | Appropriate
239 | Breast with Nodes | Image Origin | TRUE | SCN:X ±5 to ±15, Y -25 to -35, Z 6 to 12 | Appropriate
137 | Breast with Nodes | Beam Arrangement | TRUE | BRE: Tangent Pair FiF SCN: Single or Obl POP | Appropriate
238 | Breast with Nodes | Image Origin | TRUE | BRE: X & Z small, Y -25 to -45 | Appropriate
181 | Breast with Nodes | Field Shapes | TRUE |  | Appropriate
133 | Breast with Nodes | Gantry | FALSE |  | Consistent | Field Parameters
134 | Breast with Nodes | Field Length | FALSE |  | Consistent | Axial Slices
136 | Breast with Nodes | Final Collimator & Couch angles | FALSE |  | Consistent | Field Parameters
132 | Breast with Nodes | Patient Name & ID | FALSE |  | Consistent | CT Simulation Worksheet
138 | Breast with Nodes | Block Tray | FALSE | Poured Block 55 cm | Appropriate
139 | Breast with Nodes | Block Transmission | FALSE | 3.0-6MV, 3.9-15MV | Appropriate
140 | Breast with Nodes | Beam Energy | FALSE |  | Consistent | Field Parameters
141 | Breast with Nodes | Beam Weighting | FALSE | Adds up to 1.00 | Appropriate
142 | Breast with Nodes | Angles, Field Size, SSD, Wedges, Bolus | FALSE |  | Consistent | Field Parameters
143 | Breast with Nodes | Signed by Dosimetrist | FALSE |  | Correct
171 | Breast with Nodes | MUREF Location | TRUE | BRE: Near Centre of fields. SCN: At prescription point | Appropriate
180 | Breast with Nodes | Eclipse MUs | TRUE |  | Appropriate
179 | Breast with Nodes | Angles and Field Size | TRUE |  | Appropriate
135 | Breast with Nodes | Initial Collimator Angle | FALSE | Seems Reasonable | Appropriate
144 | Breast with Nodes | Plan Name & Date | FALSE |  | Consistent | TPP Plan Data Sheet
178 | Breast with Nodes | Wedges | FALSE | SCN only | Appropriate
162 | Breast with Nodes | MU_on & MU_off | FALSE |  | Appropriate
161 | Breast with Nodes | Patient & Beam Info | FALSE |  | Appropriate
149 | Breast with Nodes | Normalization | FALSE | Seems Reasonable | Appropriate
148 | Breast with Nodes | Dose coverage | FALSE | 93-109% (Prefer 95-107%) | Appropriate
147 | Breast with Nodes | Wedges | FALSE |  | Appropriate
146 | Breast with Nodes | Pencil Beam & Grid Size | FALSE | 0.6 cm | Appropriate
145 | Breast with Nodes | Unit Modeling | FALSE | UM for V3.8 Nov 2 2007 | Appropriate
29342 | Esophagus | Bolus | FALSE |  | Appropriate
29333 | Esophagus | Dose and Fractions | TRUE | 50/25 or 52.5/20 | Appropriate
29344 | Esophagus | Patient & Beam Info | FALSE |  | Appropriate
29345 | Esophagus | MU_on & MU_off | FALSE |  | Appropriate
29346 | Esophagus | User factor | FALSE |  | Appropriate
29330 | Esophagus | Diagnosis | TRUE | Esophagus | Appropriate
29332 | Esophagus | Contrast | TRUE | Contrast Present | Appropriate
29343 | Esophagus | Setup Notes | FALSE |  | Appropriate
29334 | Esophagus | Dose to Critical Organs | TRUE | Cord < 45 for 50/25, <40 for 52.5/20 | Appropriate
29335 | Esophagus | Dose to Critical Organs | TRUE | Lung: < 35% of volume receiving 20Gy (50/25) or 16Gy (52.5/20) | Appropriate
29336 | Esophagus | Dose to Critical Organs | TRUE | Heart < 50% of volume receiving 40 Gy & < 2/3 of volume receiving 45 Gy | Appropriate
29339 | Esophagus | Bolus | FALSE |  | Appropriate
29340 | Esophagus | Bolus Thickness and Extent | FALSE |  | Appropriate
29341 | Esophagus | Bolus Linked to Field | FALSE |  | Appropriate
29331 | Esophagus | Site | TRUE | ESOI | Appropriate
29337 | Esophagus | Heterogenety | TRUE | On or off depending on amount of contrast in the fields | Appropriate
29501 | Generic | Dose and Fractions | TRUE |  | Consistent | Prescription Sheet
29683 | Generic IMRT | DVO Inhomogeneity | TRUE |  | Appropriate
29681 | Generic IMRT | MU | TRUE |  | Consistent | Field Parameters
29677 | Generic IMRT | Eclipse MUs | FALSE |  | Appropriate
29676 | Generic IMRT | Dose to Critical Organs | FALSE |  | Appropriate
29673 | Generic IMRT | MUREF Location | FALSE |  | Appropriate
29670 | Generic IMRT | MLC Coverage | FALSE |  | Appropriate
29701 | Generic IMRT | Wedges | FALSE |  | Appropriate
29668 | Generic IMRT | DVO SmoothY | TRUE |  | Appropriate
29693 | Generic IMRT | Dose and Fractions | TRUE |  | Consistent | Prescription Sheet
29669 | Generic IMRT | Structure Resolution | TRUE | > 2000 Points | Appropriate
29672 | Generic IMRT | Patient & Beam Info | FALSE |  | Appropriate
29700 | Generic IMRT | Wedges | FALSE |  | Appropriate
29699 | Generic IMRT | MUREF Location | FALSE |  | Appropriate
29698 | Generic IMRT | Patient Name and ID | FALSE |  | Appropriate
29697 | Generic IMRT | Isocentre | FALSE |  | Appropriate
29696 | Generic IMRT | Image Origin | FALSE |  | Appropriate
29689 | Generic IMRT | DVO SmoothX | TRUE |  | Appropriate
29694 | Generic IMRT | SSD | FALSE |  | Appropriate
29684 | Generic IMRT | Wedges | FALSE |  | Appropriate
29690 | Generic IMRT | Field Shapes | FALSE |  | Appropriate
29667 | Generic IMRT | MU_on & MU_off | FALSE |  | Appropriate
29688 | Generic IMRT | User factor | FALSE |  | Appropriate
29687 | Generic IMRT | MU Calculation | FALSE |  | Appropriate
29686 | Generic IMRT | Dynamic Document | FALSE |  | Appropriate
29685 | Generic IMRT | 180 E | TRUE |  | Appropriate
29695 | Generic IMRT | Collimator Rotation | TRUE | Appropriate for site and gantry angle | Appropriate
29649 | Generic IMRT | Beam Energy | TRUE | 6 MV | Appropriate
29634 | Generic IMRT | Reconstructed PI dose | TRUE |  | Appropriate
29635 | Generic IMRT | Fluences | TRUE |  | Appropriate
29674 | Generic IMRT | SSD | TRUE |  | Consistent | Field Parameters
29636 | Generic IMRT | MUREF Location | FALSE |  | Appropriate
29666 | Generic IMRT | Patient Name and ID | FALSE |  | Appropriate
29640 | Generic IMRT | Isocentre | FALSE |  | Appropriate
29661 | Generic IMRT | Dose Per Field | FALSE |  | Appropriate
29648 | Generic IMRT | MLC | TRUE | Dose Dynamic | Appropriate
29631 | Generic IMRT | Isocentre Shifts | TRUE | Approximately centred on target Y1/Y2 <3 | Appropriate
29653 | Generic IMRT | Field Names | TRUE | Match with gantry angles | Appropriate
29655 | Generic IMRT | Plan Label | FALSE |  | Appropriate
29656 | Generic IMRT | Wedges | FALSE |  | Appropriate
29657 | Generic IMRT | Energy | FALSE |  | Appropriate
29658 | Generic IMRT | Isocentre | FALSE |  | Appropriate
29659 | Generic IMRT | Treatment Unit | FALSE |  | Appropriate
29660 | Generic IMRT | Effective Depth | FALSE |  | Appropriate
29646 | Generic IMRT | Isocentre | FALSE |  | Appropriate
29624 | Generic IMRT | MUs | TRUE | MU Sum agrees for Split Fields | Appropriate
29603 | Generic IMRT | SSD | FALSE |  | Appropriate
29604 | Generic IMRT | Gantry and Couch Rotation | TRUE | No couch rotation | Appropriate
29605 | Generic IMRT | Wedges | FALSE |  | Appropriate
29616 | Generic IMRT | Exported to IMSure | FALSE |  | Appropriate
29619 | Generic IMRT | Optimization Parameters | TRUE |  | Appropriate
29620 | Generic IMRT | DVO Minimize dose | TRUE |  | Appropriate
29621 | Generic IMRT | IMRT EPID Record | TRUE |  | Appropriate
29633 | Generic IMRT | Histogram | TRUE |  | Appropriate
29623 | Generic IMRT | MLC Leaf Gap | FALSE |  | Appropriate
29632 | Generic IMRT | IMRT Matrixx QA Results | TRUE |  | Appropriate
29625 | Generic IMRT | Normalization Mode | TRUE | 100% to 95% Volume | Appropriate
29626 | Generic IMRT | Bolus Linked to Field | FALSE |  | Appropriate
29627 | Generic IMRT | Optimization Minimization | TRUE |  | Appropriate
29628 | Generic IMRT | Bolus Thickness and Extent | TRUE | Included in External | Appropriate
29629 | Generic IMRT | Organs at Risk | TRUE | Correct nomencature used, Structures are smooth | Appropriate
29630 | Generic IMRT | GTV, CTV, PTV | TRUE | Correct nomenclature used, Structures are smooth | Appropriate
29639 | Generic IMRT | Original Couch Vertical | FALSE |  | Appropriate
29622 | Generic IMRT | Isodose Levels | FALSE |  | Appropriate
29664 | Generic IMRT | Patient Origin | FALSE |  | Appropriate
29665 | Generic IMRT | PSSD | FALSE |  | Appropriate
29663 | Generic IMRT | Angles and Field Size | FALSE |  | Appropriate
29662 | Generic IMRT | CAX SSD | FALSE |  | Appropriate
29024 | GYN | Dose and Fractions | TRUE | 46/23, 45/25 + 5.4/3 | Appropriate
29026 | GYN | Slice Thickness | TRUE | 5 mm | Appropriate
29027 | GYN | Plane Separation | TRUE | 0.5 cm | Appropriate
29028 | GYN | Image Registration | FALSE |  | Appropriate
29029 | GYN | Lat. Tattoo Reference | TRUE | ON AP Straightening Tattoo | Appropriate
29030 | GYN | Original Couch Vertical | FALSE |  | Appropriate
29031 | GYN | Image Origin | FALSE |  | Appropriate
29032 | GYN | MU_on & MU_off | FALSE |  | Appropriate
29033 | GYN | Patient & Beam Info | FALSE |  | Appropriate
29035 | GYN | Wedges | FALSE |  | Appropriate
29452 | GYN | Isocentre | FALSE |  | Appropriate
29037 | GYN | Wedges | FALSE |  | Appropriate
29038 | GYN | 180 E | FALSE |  | Appropriate
29039 | GYN | Calculation Volume | FALSE |  | Appropriate
29040 | GYN | PTV and CTV coverage | FALSE |  | Appropriate
29042 | GYN | Bolus | TRUE |  | Appropriate
29043 | GYN | Immobilization | TRUE | Extremity Block | Appropriate
29044 | GYN | Bolus Thickness and Extent | FALSE |  | Appropriate
29045 | GYN | Bolus Linked to Field | TRUE |  | Appropriate
29152 | GYN | Dose to Critical Organs | TRUE | Kidneys: D18Gy < 30% | Appropriate
29034 | GYN | Beam Arrangement | TRUE | 4 Field Box | Appropriate
29025 | GYN | Site | TRUE | PELB or CERV or UTER or BLAD or VULV | Appropriate
29046 | GYN | Isocentre Shifts | TRUE | No Shifts | Appropriate
29047 | GYN | Collimator Rotation | FALSE |  | Appropriate
29048 | GYN | User factor | FALSE |  | Appropriate
29049 | GYN | Portal Image Isocentre | FALSE |  | Appropriate
29151 | GYN | Organs at Risk | TRUE | Kidneys | Appropriate
29050 | GYN | Bolus | TRUE |  | Appropriate
29066 | GYN | Patient Anatomy | TRUE | Bladder / Rectum Contents | Appropriate
28920 | H&N w Comps | Energy | FALSE |  | Appropriate
28946 | H&N w Comps | MUREF | TRUE |  | Appropriate
28944 | H&N w Comps | Beam Energy | TRUE |  | Consistent | TPP Plan Data Sheet
28921 | H&N w Comps | Wedges | FALSE |  | Appropriate
28943 | H&N w Comps | Field Size | FALSE |  | Consistent | TPP Plan Data Sheet
28919 | H&N w Comps | Treatment Unit | FALSE |  | Appropriate
28918 | H&N w Comps | Isocentre | FALSE |  | Appropriate
28917 | H&N w Comps | Effective Depth | FALSE |  | Appropriate
28916 | H&N w Comps | PSSD | FALSE |  | Appropriate
28915 | H&N w Comps | Dose Per Field | FALSE |  | Appropriate
28914 | H&N w Comps | CAX SSD | FALSE |  | Appropriate
28913 | H&N w Comps | MUREF Location | FALSE |  | Appropriate
28912 | H&N w Comps | Patient Origin | FALSE |  | Appropriate
28911 | H&N w Comps | Isodose Line | FALSE |  | Appropriate
28922 | H&N w Comps | Angles and Field Size | FALSE |  | Appropriate
28933 | H&N w Comps | Dose & Fractionation | TRUE |  | Entered
28929 | H&N w Comps | Plan Label | TRUE |  | Entered
28928 | H&N w Comps | Patient Name and ID | TRUE |  | Entered
28927 | H&N w Comps | Dynamic Document | FALSE |  | Appropriate
28926 | H&N w Comps | MU Calculation | FALSE |  | Appropriate
28925 | H&N w Comps | User factor | FALSE |  | Appropriate
28924 | H&N w Comps | Field Shapes | FALSE |  | Appropriate
28923 | H&N w Comps | Eclipse MUs | FALSE |  | Appropriate
28931 | H&N w Comps | Bolus Thickness | TRUE |  | Entered
28895 | H&N w Comps | Set-up Reference Labels | TRUE |  | Consistent | Axial Slices
28945 | H&N w Comps | MU | TRUE |  | Consistent | MU Calc
28910 | H&N w Comps | Plan Label | FALSE |  | Appropriate
28930 | H&N w Comps | Plan Datasheet copied & Pasted | TRUE |  | Entered
28934 | H&N w Comps | Normalization Factor | TRUE |  | Entered
28935 | H&N w Comps | MU Calculation | TRUE |  | Correct
28936 | H&N w Comps | Calculation Printed and Signed | TRUE |  | Entered
28937 | H&N w Comps | Portal Image Isocentre | TRUE |  | Appropriate
28938 | H&N w Comps | Plan Name | TRUE |  | Consistent | TPP Dose Distribution
28939 | H&N w Comps | Treatment Machine | FALSE |  | Appropriate
28940 | H&N w Comps | Field ID | TRUE |  | Consistent | TPP Plan Data Sheet
28941 | H&N w Comps | Tolerance Table | TRUE |  | Appropriate
28942 | H&N w Comps | Gantry, Collimator, Couch angles | TRUE |  | Consistent | TPP Plan Data Sheet
28932 | H&N w Comps | Block/MLC | TRUE |  | Entered
29000 | H&N w Comps | Plane Separation | TRUE | 0.3 cm | Appropriate
28897 | H&N w Comps | Wedges | TRUE |  | Appropriate
28986 | H&N w Comps | Grid Size | FALSE |  | Appropriate
28987 | H&N w Comps | Field Normalization Type | FALSE |  | Appropriate
28988 | H&N w Comps | Heterogenety | FALSE |  | Appropriate
28989 | H&N w Comps | Calculation Model | FALSE |  | Appropriate
28993 | H&N w Comps | Hot Spots | FALSE |  | Appropriate
28994 | H&N w Comps | Isodose Levels | FALSE |  | Appropriate
28995 | H&N w Comps | MUREF Location | FALSE |  | Appropriate
28996 | H&N w Comps | Dose and Fractions | TRUE | 40/20, 10/5, 6/3 | Appropriate
28984 | H&N w Comps | Dose and Fractions | FALSE |  | Appropriate
28999 | H&N w Comps | Slice Thickness | TRUE | 3 mm | Appropriate
28983 | H&N w Comps | Dose to Critical Organs | FALSE |  | Appropriate
29001 | H&N w Comps | Contrast | FALSE |  | Appropriate
29002 | H&N w Comps | Patient Alignment | FALSE |  | Appropriate
29003 | H&N w Comps | Inhomogeneity | TRUE |  | Appropriate
29004 | H&N w Comps | Unit Modeling | TRUE |  | Appropriate
29009 | H&N w Comps | Normalization | TRUE |  | Appropriate
29008 | H&N w Comps | Isocentre Location | TRUE |  | Appropriate
29005 | H&N w Comps | Beam Energy | TRUE |  | Appropriate
29006 | H&N w Comps | Beam Weighting | TRUE |  | Appropriate
29007 | H&N w Comps | Angles, Field Size, SSD, Wedges, Bolus | TRUE |  | Appropriate
28998 | H&N w Comps | Immobilization | TRUE | Barker System | Appropriate
28898 | H&N w Comps | Dose coverage | TRUE | 93-109% (Prefer 95-107%) | Appropriate
28908 | H&N w Comps | MU_on & MU_off | FALSE |  | Appropriate
28907 | H&N w Comps | Patient & Beam Info | FALSE |  | Appropriate
28906 | H&N w Comps | Plan Approval date | FALSE |  | Appropriate
28905 | H&N w Comps | Equivalent Path Length | FALSE |  | Appropriate
28904 | H&N w Comps | Field Parameters | FALSE |  | Appropriate
28903 | H&N w Comps | ICHAT & Shifts | FALSE |  | Appropriate
28902 | H&N w Comps | Exported to IMSure | FALSE |  | Appropriate
28901 | H&N w Comps | Plan Approval | FALSE |  | Appropriate
28900 | H&N w Comps | Block Transmission | FALSE |  | Appropriate
28985 | H&N w Comps | Normalization Mode | FALSE |  | Appropriate
28899 | H&N w Comps | Block Tray | FALSE |  | Appropriate
28909 | H&N w Comps | Patient Name and ID | FALSE |  | Appropriate
28896 | H&N w Comps | Pencil Beam & Grid Size | TRUE | 0.6 cm | Appropriate
29013 | H&N w Comps | SSD | TRUE |  | Consistent | TPP Plan Data Sheet
29012 | H&N w Comps | Wedges | TRUE |  | Consistent | TPP Plan Data Sheet
29011 | H&N w Comps | Plan Name & Date | TRUE |  | Appropriate
29010 | H&N w Comps | Signed by Dosimetrist | TRUE |  | Appropriate
28978 | H&N w Comps | Straightening tattoo | FALSE |  | Appropriate
28980 | H&N w Comps | Patient Anatomy | TRUE | Dentures Out | Appropriate
28981 | H&N w Comps | Calculation Volume | FALSE |  | Appropriate
28982 | H&N w Comps | PTV and CTV coverage | FALSE |  | Appropriate
29407 | IMRT CNS | Closed Body Contour | FALSE |  | Appropriate
29589 | IMRT CNS | IMRT EPID Record | TRUE |  | Appropriate
29450 | IMRT CNS | Bolus | FALSE |  | Appropriate
29416 | IMRT CNS | Field Shapes | FALSE |  | Appropriate
29415 | IMRT CNS | User factor | FALSE |  | Appropriate
29414 | IMRT CNS | MU Calculation | FALSE |  | Appropriate
29411 | IMRT CNS | PSSD | FALSE |  | Appropriate
29410 | IMRT CNS | Bolus Thickness and Extent | FALSE |  | Appropriate
29409 | IMRT CNS | Contrast | FALSE |  | Appropriate
29418 | IMRT CNS | Angles and Field Size | FALSE |  | Appropriate
29404 | IMRT CNS | Gantry and Couch Rotation | TRUE | No couch rotation | Appropriate
29419 | IMRT CNS | Wedges | FALSE |  | Appropriate
29347 | IMRT CNS | Dose to Critical Organs | FALSE |  | Appropriate
29442 | IMRT CNS | Isocentre | FALSE |  | Appropriate
29443 | IMRT CNS | Image Origin | FALSE |  | Appropriate
29508 | IMRT CNS | MLC Leaf Gap | FALSE |  | Appropriate
29446 | IMRT CNS | Patient Name and ID | FALSE |  | Appropriate
29447 | IMRT CNS | MUREF Location | FALSE |  | Appropriate
29448 | IMRT CNS | Wedges | FALSE |  | Appropriate
29441 | IMRT CNS | Original Couch Vertical | FALSE |  | Appropriate
29408 | IMRT CNS | Scanned Region | FALSE |  | Appropriate
29428 | IMRT CNS | Patient Origin | FALSE |  | Appropriate
29440 | IMRT CNS | Lens | TRUE | 10Gy Max | Appropriate
29437 | IMRT CNS | Lat. Tattoo Reference | FALSE |  | Appropriate
29436 | IMRT CNS | Diagnosis | TRUE | Brain | Appropriate
29435 | IMRT CNS | Treatment Unit | FALSE |  | Appropriate
29433 | IMRT CNS | Exported to IMSure | FALSE |  | Appropriate
29432 | IMRT CNS | Patient & Beam Info | FALSE |  | Appropriate
29431 | IMRT CNS | MU_on & MU_off | FALSE |  | Appropriate
29417 | IMRT CNS | Eclipse MUs | FALSE |  | Appropriate
29429 | IMRT CNS | Plan Label | FALSE |  | Appropriate
29451 | IMRT CNS | Setup Notes | FALSE |  | Appropriate
29427 | IMRT CNS | MUREF Location | FALSE |  | Appropriate
29426 | IMRT CNS | CAX SSD | FALSE |  | Appropriate
29425 | IMRT CNS | Dose Per Field | FALSE |  | Appropriate
29424 | IMRT CNS | MU | TRUE |  | Consistent | Field Parameters
29423 | IMRT CNS | Effective Depth | FALSE |  | Appropriate
29422 | IMRT CNS | Field Names | TRUE | Match with gantry angles | Appropriate
29421 | IMRT CNS | Isocentre | FALSE |  | Appropriate
29420 | IMRT CNS | Energy | FALSE |  | Appropriate
29430 | IMRT CNS | Patient Name and ID | FALSE |  | Appropriate
29356 | IMRT CNS | DVO Inhomogeneity | TRUE |  | Appropriate
29449 | IMRT CNS | Wedges | FALSE |  | Appropriate
29370 | IMRT CNS | Lt & Rt Optic Nerve | TRUE | 40Gy 40/15 54Gy 60/30 | Appropriate
29369 | IMRT CNS | Chiasm | TRUE | 40Gy 40/15 54Gy 60/30 | Appropriate
29368 | IMRT CNS | Brain Stem | TRUE | 40Gy 40/15 54Gy 60/30 | Appropriate
29367 | IMRT CNS | Cord | TRUE | 40Gy 40/15 48Gy 60/30 | Appropriate
29365 | IMRT CNS | Dynamic Document | FALSE |  | Appropriate
29364 | IMRT CNS | DVO SmoothX | TRUE |  | Appropriate
29373 | IMRT CNS | Immobilization | TRUE | S Frame | Appropriate
29362 | IMRT CNS | Organs at Risk | TRUE | Correct nomencature used, Structures are smooth | Appropriate
29375 | IMRT CNS | Optimization Parameters | TRUE |  | Appropriate
29355 | IMRT CNS | DVO SmoothY | TRUE |  | Appropriate
29354 | IMRT CNS | DVO Minimize dose | TRUE |  | Appropriate
29353 | IMRT CNS | Fluences | TRUE |  | Appropriate
29352 | IMRT CNS | Reconstructed PI dose | TRUE |  | Appropriate
29351 | IMRT CNS | Histogram | TRUE |  | Appropriate
29350 | IMRT CNS | IMRT Matrixx QA Results | TRUE |  | Appropriate
29349 | IMRT CNS | Isocentre Shifts | TRUE | Approximately centred on target Y1/Y2 <3 | Appropriate
29348 | IMRT CNS | GTV, CTV, PTV | TRUE | Correct nomenclature used, Structures are smooth, Margin to PTV is 0.5 cm and all PTVs are 5 mm from body | Appropriate
29363 | IMRT CNS | Straightening tattoo | FALSE |  | Appropriate
29389 | IMRT CNS | Wedges | FALSE |  | Appropriate
29403 | IMRT CNS | MUREF Location | FALSE |  | Appropriate
29402 | IMRT CNS | Normalization Mode | TRUE | 100% to 95% Volume | Appropriate
29401 | IMRT CNS | MUs | TRUE | MU Sum agrees for Split Fields | Appropriate
29399 | IMRT CNS | Isodose Levels | FALSE |  | Appropriate
29397 | IMRT CNS | Bolus Linked to Field | FALSE |  | Appropriate
29396 | IMRT CNS | MLC Coverage | FALSE |  | Appropriate
29395 | IMRT CNS | Structure Resolution | TRUE | > 2000 Points | Appropriate
29371 | IMRT CNS | Lt & Rt Orbit | TRUE | 40Gy 40/15 45Gy 60/30 | Appropriate
29390 | IMRT CNS | MLC | TRUE | Dose Dynamic | Appropriate
29405 | IMRT CNS | SSD | FALSE |  | Appropriate
29388 | IMRT CNS | Collimator Rotation | FALSE |  | Appropriate
29387 | IMRT CNS | Field Jaw Settings | FALSE |  | Appropriate
29386 | IMRT CNS | Wedges | FALSE |  | Appropriate
29385 | IMRT CNS | Hot Spots | FALSE |  | Appropriate
29381 | IMRT CNS | Dose and Fractions | TRUE | 60/30, 40/15 | Appropriate
29380 | IMRT CNS | Patient Alignment | FALSE |  | Appropriate
29379 | IMRT CNS | Plane Separation | TRUE | 0.3 cm | Appropriate
29378 | IMRT CNS | Slice Thickness | TRUE | 3 mm | Appropriate
29391 | IMRT CNS | Beam Energy | TRUE | 6 MV | Appropriate
29067 | IMRT Head & Neck | Contrast | FALSE |  | Appropriate
29069 | IMRT Head & Neck | Scanned Region | FALSE |  | Appropriate
29071 | IMRT Head & Neck | Beam Arrangement | TRUE | 150, 90, 45, 0, 310, 250, 200 | Appropriate
29022 | IMRT Head & Neck | Dose and Fractions | TRUE | 70/35, 66/33, 60/30 | Appropriate
29072 | IMRT Head & Neck | SSD | FALSE |  | Appropriate
29073 | IMRT Head & Neck | Gantry and Couch Rotation | TRUE | No couch rotation | Appropriate
29074 | IMRT Head & Neck | Wedges | FALSE |  | Appropriate
29075 | IMRT Head & Neck | Collimator Rotation | FALSE | Post-lat fields (within 45 degrees of horizontal) rotated to come off shoulders | Appropriate
29070 | IMRT Head & Neck | Closed Body Contour | FALSE |  | Appropriate
29023 | IMRT Head & Neck | Patient Anatomy | TRUE | Dentures Out | Appropriate
29021 | IMRT Head & Neck | Patient Alignment | FALSE |  | Appropriate
29020 | IMRT Head & Neck | Plane Separation | TRUE | 0.3 cm | Appropriate
29019 | IMRT Head & Neck | Slice Thickness | TRUE | 3 mm | Appropriate
29018 | IMRT Head & Neck | Immobilization | TRUE | S Frame | Appropriate
29016 | IMRT Head & Neck | Straightening tattoo | FALSE |  | Appropriate
29015 | IMRT Head & Neck | Contrast | FALSE |  | Appropriate
29014 | IMRT Head & Neck | Contrast | FALSE |  | Appropriate
29076 | IMRT Head & Neck | Field Jaw Settings | TRUE | Post-lat fields (within 45 degrees of horizontal) adjusted to come off shoulders | Appropriate
29117 | IMRT Head & Neck | DVO SmoothY | TRUE |  | Appropriate
29126 | IMRT Head & Neck | Optimization Parameters | TRUE |  | Appropriate
29118 | IMRT Head & Neck | DVO Minimize dose | TRUE |  | Appropriate
29590 | IMRT Head & Neck | IMRT EPID Record | TRUE |  | Appropriate
29325 | IMRT Head & Neck | Isodose Levels | FALSE |  | Appropriate
29326 | IMRT Head & Neck | MLC Leaf Gap | FALSE |  | Appropriate
29327 | IMRT Head & Neck | MUs | TRUE | MU Sum agrees for Split Fields | Appropriate
29328 | IMRT Head & Neck | Normalization Mode | TRUE | 100% to 95% Volume | Appropriate
29323 | IMRT Head & Neck | Bolus Linked to Field | FALSE |  | Appropriate
29127 | IMRT Head & Neck | Optimization Minimization | FALSE |  | Appropriate
29322 | IMRT Head & Neck | Bolus Thickness and Extent | TRUE | Included in External | Appropriate
29125 | IMRT Head & Neck | Organs at Risk | TRUE | Correct nomencature used, Structures are smooth | Appropriate
29124 | IMRT Head & Neck | GTV, CTV, PTV | TRUE | Correct nomenclature used, Structures are smooth, Margin to PTV is 0.5 cm | Appropriate
29123 | IMRT Head & Neck | Isocentre Shifts | TRUE | Approximately centred on target Y1/Y2 <3 | Appropriate
29122 | IMRT Head & Neck | IMRT Matrixx QA Results | TRUE |  | Appropriate
29121 | IMRT Head & Neck | Histogram | TRUE |  | Appropriate
29120 | IMRT Head & Neck | Reconstructed PI dose | TRUE |  | Appropriate
29119 | IMRT Head & Neck | Fluences | TRUE |  | Appropriate
29329 | IMRT Head & Neck | MUREF Location | FALSE |  | Appropriate
29161 | IMRT Head & Neck | Cord | TRUE | V45Gy ≤ 0.1cc, V48Gy ≤ 0cc | Appropriate
29438 | IMRT Head & Neck | Lat. Tattoo Reference | FALSE |  | Appropriate
29115 | IMRT Head & Neck | DVO Inhomogeneity | TRUE |  | Appropriate
29500 | IMRT Head & Neck | Isocentre | FALSE |  | Appropriate
29166 | IMRT Head & Neck | Larynx | TRUE | V50Gy ≤ 67% Mean Dose ≤ 45Gy | Appropriate
29165 | IMRT Head & Neck | Lt & Rt Orbit | TRUE | V45Gy ≤ 0% | Appropriate
29164 | IMRT Head & Neck | Lt & Rt Optic Nerve | TRUE | V50Gy ≤ 0% | Appropriate
29324 | IMRT Head & Neck | Hot Spots | FALSE |  | Appropriate
29162 | IMRT Head & Neck | Brain Stem | TRUE | V50Gy ≤ 0.1cc, V54Gy ≤ 0cc | Appropriate
29444 | IMRT Head & Neck | Isocentre | FALSE |  | Appropriate
29160 | IMRT Head & Neck | Mandible | TRUE | V70Gy ≤ 0.1cc | Appropriate
29313 | IMRT Head & Neck | MLC | TRUE | Dose Dynamic | Appropriate
29314 | IMRT Head & Neck | Beam Energy | TRUE | 6 MV | Appropriate
29317 | IMRT Head & Neck | Elective Nodes (PTV56) | TRUE | D99% ≥ 53.2 Gy, D95% ≥ 56 Gy V77Gy ≤ 0% | Appropriate
29318 | IMRT Head & Neck | Parotid & Submandibulars | TRUE | Rt or Lt V30Gy ≤ 50% Rt or Lt Mean Dose ≤ 26Gy Combined Rt & Lt V20Gy ≤ 20cc | Appropriate
29319 | IMRT Head & Neck | Brachial Plexus | TRUE | V60Gy < 10% | Appropriate
29078 | IMRT Head & Neck | Field Names | TRUE | Match with gantry angles | Appropriate
29163 | IMRT Head & Neck | Chiasm | TRUE | V50Gy ≤ 0% | Appropriate
29084 | IMRT Head & Neck | MU_on & MU_off | FALSE |  | Appropriate
29094 | IMRT Head & Neck | Isocentre | FALSE |  | Appropriate
29093 | IMRT Head & Neck | Treatment Unit | FALSE |  | Appropriate
29092 | IMRT Head & Neck | Effective Depth | FALSE |  | Appropriate
29091 | IMRT Head & Neck | PSSD | FALSE |  | Appropriate
29090 | IMRT Head & Neck | Dose Per Field | FALSE |  | Appropriate
29088 | IMRT Head & Neck | MUREF Location | FALSE |  | Appropriate
29087 | IMRT Head & Neck | Patient Origin | FALSE |  | Appropriate
29095 | IMRT Head & Neck | Energy | FALSE |  | Appropriate
29085 | IMRT Head & Neck | Patient Name and ID | FALSE |  | Appropriate
29089 | IMRT Head & Neck | CAX SSD | FALSE |  | Appropriate
29083 | IMRT Head & Neck | Patient & Beam Info | FALSE |  | Appropriate
29082 | IMRT Head & Neck | Exported to IMSure | FALSE |  | Appropriate
29116 | IMRT Head & Neck | DVO SmoothX | TRUE |  | Appropriate
29321 | IMRT Head & Neck | Structure Resolution | TRUE | > 2000 Points | Appropriate
29079 | IMRT Head & Neck | MLC Coverage | FALSE |  | Appropriate
29080 | IMRT Head & Neck | Original Couch Vertical | FALSE |  | Appropriate
29081 | IMRT Head & Neck | Image Origin | FALSE |  | Appropriate
29086 | IMRT Head & Neck | Plan Label | FALSE |  | Appropriate
29103 | IMRT Head & Neck | 180 E | FALSE |  | Appropriate
29111 | IMRT Head & Neck | Intermediate Risk (PTV60) | TRUE | D99% ≥ 55.8 Gy, D95% ≥ 60 Gy V72.6Gy ≤ 0%-Resiual Disease V66Gy ≤ 2%,V69Gy ≤ 0%-No Resiual Disease | Appropriate
29112 | IMRT Head & Neck | Elective Nodes (PTV54) | TRUE | D99% ≥ 50.2 Gy, D95% ≥ 54 Gy V72.6Gy ≤ 0% | Appropriate
29096 | IMRT Head & Neck | Wedges | FALSE |  | Appropriate
29109 | IMRT Head & Neck | Residual Disease (PTV66) | TRUE | D99% ≥ 61.4 Gy, D95% ≥ 66 Gy V72.6Gy ≤ 2% V75.9Gy ≤0% | Appropriate
29108 | IMRT Head & Neck | Gross Disease (PTV70) | TRUE | D99% ≥ 65.1 Gy, D95% ≥ 70 Gy V77Gy ≤ 2% V80.5Gy ≤0% | Appropriate
29106 | IMRT Head & Neck | MU | TRUE |  | Consistent | Field Parameters
29104 | IMRT Head & Neck | Wedges | FALSE |  | Appropriate
29110 | IMRT Head & Neck | Intermediate Risk (PTV63) | TRUE | D99% ≥ 58.6 Gy, D95% ≥ 63 Gy V77Gy ≤ 0% | Appropriate
29113 | IMRT Head & Neck | PTV and CTV coverage | FALSE |  | Appropriate
29102 | IMRT Head & Neck | Dynamic Document | FALSE |  | Appropriate
29101 | IMRT Head & Neck | MU Calculation | FALSE |  | Appropriate
29100 | IMRT Head & Neck | User factor | FALSE |  | Appropriate
29099 | IMRT Head & Neck | Field Shapes | FALSE |  | Appropriate
29098 | IMRT Head & Neck | Eclipse MUs | FALSE |  | Appropriate
29114 | IMRT Head & Neck | Dose to Critical Organs | FALSE |  | Appropriate
29097 | IMRT Head & Neck | Angles and Field Size | FALSE |  | Appropriate
29105 | IMRT Head & Neck | SSD | TRUE |  | Consistent | Field Parameters
29289 | IMRT Prostate | PSSD | FALSE |  | Appropriate
29284 | IMRT Prostate | Plan Label | FALSE |  | Appropriate
29288 | IMRT Prostate | Dose Per Field | FALSE |  | Appropriate
29287 | IMRT Prostate | CAX SSD | FALSE |  | Appropriate
29290 | IMRT Prostate | Effective Depth | FALSE |  | Appropriate
29286 | IMRT Prostate | MUREF Location | FALSE |  | Appropriate
29285 | IMRT Prostate | Patient Origin | FALSE |  | Appropriate
29283 | IMRT Prostate | Patient Name and ID | FALSE |  | Appropriate
29282 | IMRT Prostate | MU_on & MU_off | FALSE |  | Appropriate
29281 | IMRT Prostate | Patient & Beam Info | FALSE |  | Appropriate
29280 | IMRT Prostate | Exported to IMSure | FALSE |  | Appropriate
29279 | IMRT Prostate | Image Origin | FALSE |  | Appropriate
29276 | IMRT Prostate | MLC Coverage | FALSE |  | Appropriate
29277 | IMRT Prostate | Fluences | TRUE |  | Appropriate
29291 | IMRT Prostate | Isocentre | FALSE |  | Appropriate
29309 | IMRT Prostate | Histogram | TRUE |  | Consistent | Field Parameters
29275 | IMRT Prostate | MUs | FALSE |  | Appropriate
29274 | IMRT Prostate | Field Jaw Settings | FALSE |  | Appropriate
29278 | IMRT Prostate | Original Couch Vertical | FALSE |  | Appropriate
29302 | IMRT Prostate | Bolus | FALSE |  | Appropriate
29507 | IMRT Prostate | MLC Leaf Gap | FALSE |  | Appropriate
29273 | IMRT Prostate | Wedges | FALSE |  | Appropriate
29312 | IMRT Prostate | MLC | TRUE | Dose Dynamic | Appropriate
29311 | IMRT Prostate | Equivalent Path Length | FALSE |  | Appropriate
29310 | IMRT Prostate | IMRT Matrixx QA Results | TRUE |  | Appropriate
29308 | IMRT Prostate | Reconstructed PI dose | TRUE |  | Consistent | Field Parameters
29306 | IMRT Prostate | SSD | FALSE |  | Appropriate
29305 | IMRT Prostate | Wedges | FALSE |  | Appropriate
29307 | IMRT Prostate | MU | TRUE |  | Appropriate
29303 | IMRT Prostate | Setup Notes | FALSE |  | Appropriate
29292 | IMRT Prostate | Treatment Unit | FALSE |  | Appropriate
29301 | IMRT Prostate | Portal Image Isocentre | FALSE |  | Appropriate
29300 | IMRT Prostate | Dynamic Document | FALSE |  | Appropriate
29299 | IMRT Prostate | MU Calculation | FALSE |  | Appropriate
29298 | IMRT Prostate | User factor | FALSE |  | Appropriate
29297 | IMRT Prostate | Field Shapes | FALSE |  | Appropriate
29296 | IMRT Prostate | Eclipse MUs | FALSE |  | Appropriate
29295 | IMRT Prostate | Angles and Field Size | FALSE |  | Appropriate
29294 | IMRT Prostate | Wedges | FALSE |  | Appropriate
29293 | IMRT Prostate | Energy | FALSE |  | Appropriate
29304 | IMRT Prostate | 180 E | FALSE |  | Appropriate
29246 | IMRT Prostate | Plane Separation | TRUE | 0.3 cm | Appropriate
29506 | IMRT Prostate | Isocentre | FALSE |  | Appropriate
29254 | IMRT Prostate | Dose to Critical Organs | FALSE |  | Appropriate
29253 | IMRT Prostate | (CTV70) | TRUE | D100% ≥ 70 Gy | Appropriate
29252 | IMRT Prostate | (CTV76) | TRUE | D100% ≥ 76 Gy | Appropriate
29251 | IMRT Prostate | Gross Disease (PTV76) | TRUE | D98% ≥ 76 Gy V81Gy ≤ 2% V83Gy ≤0% | Appropriate
29250 | IMRT Prostate | Gross Disease (PTV70) | TRUE | D98% ≥ 70 Gy V75Gy ≤ 2% V77Gy ≤0% | Appropriate
29249 | IMRT Prostate | PTV and CTV coverage | FALSE |  | Appropriate
29505 | IMRT Prostate | Immobilization | TRUE | Extremity Block | Appropriate
29272 | IMRT Prostate | Collimator Rotation | FALSE |  | Appropriate
29247 | IMRT Prostate | GTV, CTV, PTV | TRUE | Correct nomenclature used, Structures are smooth | Appropriate
29244 | IMRT Prostate | Lat. Tattoo Reference | TRUE | At Level of AP tattoo | Appropriate
29243 | IMRT Prostate | Image Registration | FALSE |  | Appropriate
29242 | IMRT Prostate | Patient Anatomy | TRUE | Bladder Full, Rectum Empty | Appropriate
29241 | IMRT Prostate | Dose and Fractions | TRUE | 70/35 (low Risk), 76/38 (Int. Risk), 46/23 & 24/12 | Appropriate
29240 | IMRT Prostate | Bolus | FALSE |  | Appropriate
29239 | IMRT Prostate | Site | TRUE | PELB, PROS | Appropriate
29445 | IMRT Prostate | Isocentre | FALSE |  | Appropriate
29591 | IMRT Prostate | IMRT EPID Record | TRUE |  | Appropriate
29248 | IMRT Prostate | Organs at Risk | TRUE | Bladder, Rectum, Femural Heads: Correct nomenclature used, Structures are smooth | Appropriate
29263 | IMRT Prostate | DVO Minimize dose | TRUE |  | Appropriate
29245 | IMRT Prostate | Slice Thickness | TRUE | 3 mm | Appropriate
29271 | IMRT Prostate | Beam Energy | TRUE | 6 MV | Appropriate
29503 | IMRT Prostate | MUREF Location | FALSE |  | Appropriate
29269 | IMRT Prostate | Bolus Thickness and Extent | FALSE |  | Appropriate
29268 | IMRT Prostate | Gantry and Couch Rotation | TRUE | No couch rotation | Appropriate
29267 | IMRT Prostate | Field Names | TRUE | Match with gantry angles | Appropriate
29265 | IMRT Prostate | Isocentre Shifts | FALSE | Approximately centred on target Y1/Y2 <3 | Appropriate
29270 | IMRT Prostate | Bolus Linked to Field | FALSE |  | Appropriate
29264 | IMRT Prostate | Beam Arrangement | TRUE | 110,80,40,355,310,280,250 | Appropriate
29262 | IMRT Prostate | DVO SmoothY | TRUE | 40 | Appropriate
29255 | IMRT Prostate | Bladder | TRUE | V80Gy ≤ 15%, V75Gy ≤ 25%, V70Gy ≤ 35%, V50Gy ≤ 65% | Appropriate
29260 | IMRT Prostate | DVO Inhomogeneity | TRUE |  | Appropriate
29259 | IMRT Prostate | Optimization Minimization | FALSE |  | Appropriate
29258 | IMRT Prostate | Optimization Parameters | TRUE |  | Appropriate
29257 | IMRT Prostate | Femural Head | TRUE | V52Gy ≤ 0% | Appropriate
29256 | IMRT Prostate | Rectum | TRUE | V75Gy ≤ 15%, V70Gy ≤ 25%, V65Gy ≤ 35%, V60Gy ≤ 50% | Appropriate
29261 | IMRT Prostate | DVO SmoothX | TRUE | 50 | Appropriate
29502 | IMRT Prostate | Normalization Mode | TRUE | 100% covers 98% vol. | Appropriate
11 | Lung | Plane Separation | TRUE | 0.5 cm | Appropriate
10 | Lung | Slice Thickness | TRUE | 5 mm | Appropriate
8 | Lung | Dose and Fractions | TRUE | 40/15 (small cell), 60/15, 40/20 & 22/10 or 55/20 (NSCLC), 39/13 or 30/10 (Palliative) | Appropriate
29467 | Lung | Bolus Thickness and Extent | FALSE |  | Appropriate
29128 | Lung | Bolus | FALSE |  | Appropriate
29130 | Lung | Dose to Critical Organs | TRUE | Heart: 2/3 Vol. Gets < 45 Gy | Appropriate
29129 | Lung | Dose to Critical Organs | TRUE | Cord: <45 Gy or 35Gy (Small Cell) | Appropriate
21 | Lung | Dose to Critical Organs | TRUE | Lung: <35% vol. Gets > 20Gy or 16Gy (55/20), <30% vol. Gets > 10Gy (60/15) | Appropriate
29137 | Lung | Image Origin | TRUE | X: -R +L, Y -25 to -35, Z small | Appropriate
130 | Lung | Beam Energy | TRUE | Not 15 MV alone, 6MV for f/s < 10x10 | Appropriate
16 | Lung | Bolus | FALSE |  | Appropriate
17 | Lung | Immobilization | TRUE | Double MBS | Appropriate
18 | Lung | Organs at Risk | TRUE | Lung, Heart, Cord | Appropriate
20 | Lung | Patient Anatomy | FALSE |  | Appropriate
90 | Lung | PTV and CTV coverage | TRUE | PTV gets 90% CTV gets 95%, PTV gets 95% for 60/15 | Appropriate
28972 | Lung | Site | TRUE | LUN L or R | Appropriate
29466 | Lung | Bolus Linked to Field | FALSE |  | Appropriate
19 | Lung | Setup Notes | FALSE |  | Appropriate
29463 | Pancreas | Organs at Risk | TRUE | Kidneys, Liver, Small Bowel, Cord | Appropriate
29462 | Pancreas | Setup Notes | FALSE |  | Appropriate
29461 | Pancreas | Patient Anatomy | FALSE |  | Appropriate
29459 | Pancreas | Image Origin | TRUE | X: -5 to 5, Y -25 to -35, Z small | Appropriate
29458 | Pancreas | Dose to Critical Organs | TRUE | Kidneys: 2/3 Vol. Gets < 18 Gy | Appropriate
29455 | Pancreas | Immobilization | TRUE | Double MBS | Appropriate
29456 | Pancreas | Bolus | FALSE |  | Appropriate
29457 | Pancreas | Dose to Critical Organs | TRUE | Cord: <40 Gy | Appropriate
29468 | Pancreas | Site | TRUE | PANC | Appropriate
29469 | Pancreas | Dose and Fractions | TRUE | 45/25, 40/20 x2 , or 20/10 (Palliative) | Appropriate
29470 | Pancreas | Bolus | FALSE |  | Appropriate
29471 | Pancreas | Slice Thickness | TRUE | 5 mm | Appropriate
29472 | Pancreas | Dose to Critical Organs | TRUE | Liver: 1/2 Vol. Gets < 25 Gy | Appropriate
29473 | Pancreas | Dose to Critical Organs | TRUE | Small Bowel: < 2l >45 Gy | Appropriate
29476 | Pancreas | Contrast | TRUE | In Small Bowel | Appropriate
29477 | Pancreas | Contrast | TRUE | In Small Bowel | Appropriate
29479 | Pancreas | Collimator Rotation | TRUE | 270 for Dynamic Wedge | Appropriate
15 | Pancreas | Bolus Linked to Field | FALSE |  | Appropriate
14 | Pancreas | Bolus Thickness and Extent | FALSE |  | Appropriate
29478 | Pancreas | Image Registration | FALSE |  | Appropriate
276 | Prostate 2 Phase | Patient Anatomy | TRUE | Bladder Full, Rectum Empty | Appropriate
29141 | Prostate 2 Phase | Image Registration | FALSE |  | Appropriate
287 | Prostate 2 Phase | Slice Thickness | TRUE | 3 mm | Appropriate
288 | Prostate 2 Phase | Plane Separation | TRUE | 0.3 cm | Appropriate
289 | Prostate 2 Phase | Lat. Tattoo Reference | TRUE | At Level of AP tattoo | Appropriate
290 | Prostate 2 Phase | Site | TRUE | PELB, PROS | Appropriate
291 | Prostate 2 Phase | 180 E | FALSE |  | Appropriate
286 | Prostate 2 Phase | Bolus Linked to Field | FALSE |  | Appropriate
307 | Prostate 2 Phase | Bolus | FALSE |  | Appropriate
28969 | Prostate 2 Phase | Original Couch Vertical | FALSE |  | Appropriate
29143 | Prostate 2 Phase | Patient & Beam Info | FALSE |  | Appropriate
28968 | Prostate 2 Phase | MU_on & MU_off | FALSE |  | Appropriate
29150 | Prostate 2 Phase | Dose to Critical Organs | TRUE | Femural Heads < 100cc Vol. Gets 60 Gy | Appropriate
28970 | Prostate 2 Phase | Image Origin | FALSE |  | Appropriate
28975 | Prostate 2 Phase | User factor | FALSE |  | Appropriate
29149 | Prostate 2 Phase | Dose to Critical Organs | TRUE | Bladder <25% Vol. Gets > 70 Gy | Appropriate
306 | Prostate 2 Phase | Portal Image Isocentre | FALSE |  | Appropriate
275 | Prostate 2 Phase | Organs at Risk | TRUE | Bladder & 11 cm Rectum Contoured | Appropriate
285 | Prostate 2 Phase | Dose to Critical Organs | TRUE | Rectum < 50% Vol. Gets > 60 Gy | Appropriate
274 | Prostate 2 Phase | GTV, CTV, PTV | TRUE | Prostate Shape looks reasonable | Appropriate
272 | Prostate 2 Phase | MLC Coverage | TRUE | 1.5 cm margin around prostate | Appropriate
273 | Prostate 2 Phase | Dose and Fractions | TRUE | 46/23 (PELB), 24/12 (PROS) or 20/10 (post-Op) | Appropriate
29139 | Prostate 2 Phase | Bolus | FALSE |  | Appropriate
277 | Prostate 2 Phase | Bolus Thickness and Extent | FALSE |  | Appropriate
279 | Prostate 2 Phase | Setup Notes | FALSE |  | Appropriate
281 | Prostate 2 Phase | Beam Arrangement | TRUE | 4 field Box | Appropriate
282 | Prostate 2 Phase | Beam Energy | TRUE | 15 MV | Appropriate
283 | Prostate 2 Phase | Collimator Rotation | FALSE |  | Appropriate
284 | Prostate 2 Phase | PTV and CTV coverage | FALSE |  | Appropriate
29560 | Prostate 2 Phase with IMRT | Equivalent Path Length | FALSE |  | Appropriate
29551 | Prostate 2 Phase with IMRT | DVO Minimize dose | TRUE |  | Appropriate
29555 | Prostate 2 Phase with IMRT | Site | TRUE | PELB, PROS | Appropriate
29554 | Prostate 2 Phase with IMRT | Rectum | TRUE | V75Gy ≤ 15%, V70Gy ≤ 25%, V65Gy ≤ 35%, V60Gy ≤ 50% | Appropriate
29553 | Prostate 2 Phase with IMRT | Isocentre Shifts | FALSE | Approximately centred on target Y1/Y2 <3 | Appropriate
29552 | Prostate 2 Phase with IMRT | Beam Arrangement | TRUE | Phase 1: 4 field Box; Phase II: 110,80,40,355,310,280,250 | Appropriate
29559 | Prostate 2 Phase with IMRT | MLC | TRUE | Phase I: Static; Phase II: Dose Dynamic | Appropriate
29550 | Prostate 2 Phase with IMRT | DVO SmoothY | TRUE | 40 | Appropriate
29561 | Prostate 2 Phase with IMRT | IMRT Matrixx QA Results | TRUE |  | Appropriate
29549 | Prostate 2 Phase with IMRT | DVO SmoothX | TRUE | 50 | Appropriate
29584 | Prostate 2 Phase with IMRT | Bolus Thickness and Extent | FALSE |  | Appropriate
29547 | Prostate 2 Phase with IMRT | Optimization Minimization | FALSE |  | Appropriate
29572 | Prostate 2 Phase with IMRT | Plane Separation | TRUE | 0.3 cm | Appropriate
29546 | Prostate 2 Phase with IMRT | GTV, CTV, PTV | TRUE | Correct nomenclature used, Structures are smooth | Appropriate
29545 | Prostate 2 Phase with IMRT | Femural Head | TRUE | V52Gy ≤ 0% | Appropriate
29544 | Prostate 2 Phase with IMRT | Organs at Risk | TRUE | Bladder, 11 cm Rectum, Femural Heads: Correct nomenclature used, Structures are smooth | Appropriate
29548 | Prostate 2 Phase with IMRT | DVO Inhomogeneity | TRUE |  | Appropriate
29569 | Prostate 2 Phase with IMRT | Image Registration | FALSE |  | Appropriate
29565 | Prostate 2 Phase with IMRT | Bolus | FALSE |  | Appropriate
29566 | Prostate 2 Phase with IMRT | Dose and Fractions | TRUE | 46/23 (PELB), 24/12 (PROS) or 20/10 (post-Op) | Appropriate
29567 | Prostate 2 Phase with IMRT | Patient Anatomy | TRUE | Bladder Full, Rectum Empty | Appropriate
29564 | Prostate 2 Phase with IMRT | Patient Name and ID | FALSE |  | Appropriate
29563 | Prostate 2 Phase with IMRT | PTV and CTV coverage | FALSE |  | Appropriate
29543 | Prostate 2 Phase with IMRT | Bladder | TRUE | V80Gy ≤ 15%, V75Gy ≤ 25%, V70Gy ≤ 35%, V50Gy ≤ 65% | Appropriate
29575 | Prostate 2 Phase with IMRT | Collimator Rotation | FALSE |  | Appropriate
29568 | Prostate 2 Phase with IMRT | Patient & Beam Info | FALSE |  | Appropriate
29582 | Prostate 2 Phase with IMRT | Original Couch Vertical | FALSE |  | Appropriate
29570 | Prostate 2 Phase with IMRT | Lat. Tattoo Reference | TRUE | At Level of AP tattoo | Appropriate
29571 | Prostate 2 Phase with IMRT | Slice Thickness | TRUE | 3 mm | Appropriate
29574 | Prostate 2 Phase with IMRT | MU_on & MU_off | FALSE |  | Appropriate
29577 | Prostate 2 Phase with IMRT | Field Jaw Settings | TRUE |  | Appropriate
29579 | Prostate 2 Phase with IMRT | MLC Coverage | FALSE |  | Appropriate
29580 | Prostate 2 Phase with IMRT | Fluences | TRUE |  | Appropriate
29581 | Prostate 2 Phase with IMRT | Gantry and Couch Rotation | TRUE | Phase II: No couch rotation | Appropriate
29562 | Prostate 2 Phase with IMRT | Histogram | TRUE |  | Consistent | Field Parameters
29519 | Prostate 2 Phase with IMRT | MU | TRUE |  | Appropriate
29594 | Prostate 2 Phase with IMRT | Bolus | FALSE |  | Appropriate
29593 | Prostate 2 Phase with IMRT | Image Origin | FALSE |  | Appropriate
29592 | Prostate 2 Phase with IMRT | IMRT EPID Record | TRUE |  | Appropriate
29586 | Prostate 2 Phase with IMRT | Beam Energy | TRUE | Phase I: 15 MV; Phase II: 6 MV | Appropriate
29516 | Prostate 2 Phase with IMRT | 180 E | FALSE |  | Appropriate
29542 | Prostate 2 Phase with IMRT | Dose to Critical Organs | FALSE |  | Appropriate
29520 | Prostate 2 Phase with IMRT | Setup Notes | FALSE |  | Appropriate
29522 | Prostate 2 Phase with IMRT | Wedges | TRUE | Phase 1: Sup-Inf Wedge as needed | Appropriate
29524 | Prostate 2 Phase with IMRT | Isocentre | FALSE |  | Appropriate
29526 | Prostate 2 Phase with IMRT | MLC Leaf Gap | TRUE | Minimized by Jaw Settings | Appropriate
29528 | Prostate 2 Phase with IMRT | Reconstructed PI dose | TRUE |  | Consistent | Field Parameters
29529 | Prostate 2 Phase with IMRT | Isocentre | FALSE |  | Appropriate
29531 | Prostate 2 Phase with IMRT | Normalization Mode | TRUE | Phase II: 100% covers 98% vol. | Appropriate
29533 | Prostate 2 Phase with IMRT | Immobilization | TRUE | Extremity Block | Appropriate
29585 | Prostate 2 Phase with IMRT | Bolus Linked to Field | FALSE |  | Appropriate
29541 | Prostate 2 Phase with IMRT | (CTV70) | TRUE | D100% ≥ 70 Gy | Appropriate
29514 | Prostate 2 Phase with IMRT | Bolus | FALSE |  | Appropriate
29540 | Prostate 2 Phase with IMRT | (CTV76) | FALSE | D100% ≥ 76 Gy | Appropriate
29534 | Prostate 2 Phase with IMRT | Isocentre | FALSE |  | Appropriate
29588 | Prostate 2 Phase with IMRT | Isocentre | FALSE |  | Appropriate
29537 | Prostate 2 Phase with IMRT | Field Names | TRUE | Match with gantry angles | Appropriate
29538 | Prostate 2 Phase with IMRT | Gross Disease (PTV70) | TRUE | D98% ≥ 70 Gy V75Gy ≤ 2% V77Gy ≤0% | Appropriate
29539 | Prostate 2 Phase with IMRT | Gross Disease (PTV76) | FALSE | D98% ≥ 76 Gy V81Gy ≤ 2% V83Gy ≤0% | Appropriate
29536 | Prostate 2 Phase with IMRT | Optimization Parameters | TRUE |  | Appropriate
101 | Prostate Single Phase | Setup Notes | FALSE |  | Appropriate
105 | Prostate Single Phase | Collimator Rotation | FALSE |  | Appropriate
104 | Prostate Single Phase | Beam Energy | TRUE | 15 MV | Appropriate
106 | Prostate Single Phase | PTV and CTV coverage | FALSE |  | Appropriate
102 | Prostate Single Phase | Wedges | FALSE |  | Appropriate
100 | Prostate Single Phase | Bolus | FALSE |  | Appropriate
99 | Prostate Single Phase | Bolus Linked to Field | FALSE |  | Appropriate
98 | Prostate Single Phase | Bolus Thickness and Extent | FALSE |  | Appropriate
97 | Prostate Single Phase | Patient Anatomy | TRUE | Bladder Full, Rectum Empty | Appropriate
96 | Prostate Single Phase | Organs at Risk | TRUE | Bladder & Rectum Contoured | Appropriate
95 | Prostate Single Phase | GTV, CTV, PTV | TRUE | Prostate Shape looks reasonable | Appropriate
103 | Prostate Single Phase | Beam Arrangement | TRUE | 4 field Box | Appropriate
29140 | Prostate Single Phase | Bolus | FALSE |  | Appropriate
108 | Prostate Single Phase | MLC Coverage | TRUE | 1.5 cm margin around prostate | Appropriate
29146 | Prostate Single Phase | MU_on & MU_off | FALSE |  | Appropriate
292 | Prostate Single Phase | 180 E | FALSE |  | Appropriate
29142 | Prostate Single Phase | Image Registration | FALSE |  | Appropriate
28971 | Prostate Single Phase | Site | TRUE | PROS | Appropriate
28976 | Prostate Single Phase | User factor | FALSE |  | Appropriate
94 | Prostate Single Phase | Dose and Fractions | TRUE | 70/35 (low Risk), 76/38 (Int. Risk), 66/33 (Post Op) | Appropriate
29145 | Prostate Single Phase | Patient & Beam Info | FALSE |  | Appropriate
29148 | Prostate Single Phase | Dose to Critical Organs | TRUE | Femural Heads < 100cc Vol. Gets 60 Gy | Appropriate
29147 | Prostate Single Phase | Dose to Critical Organs | TRUE | Bladder <25% Vol. Gets > 70 Gy | Appropriate
107 | Prostate Single Phase | Dose to Critical Organs | TRUE | Rectum < 50% Vol. Gets > 60 Gy | Appropriate
7 | Rectum/Bladder/Anal cannal | Contrast | TRUE | Contrast in Rectum or Bladder | Appropriate
119 | Rectum/Bladder/Anal cannal | Bolus Thickness and Extent | FALSE |  | Appropriate
28977 | Rectum/Bladder/Anal cannal | User factor | FALSE |  | Appropriate
115 | Rectum/Bladder/Anal cannal | Plane Separation | TRUE | 0.5 cm | Appropriate
116 | Rectum/Bladder/Anal cannal | Offset from DICOM origin | TRUE | X, Y signs flipped from Image Origin | Appropriate
120 | Rectum/Bladder/Anal cannal | Bolus Linked to Field | FALSE |  | Appropriate
117 | Rectum/Bladder/Anal cannal | Heterogenety | TRUE | OFF (due to contrast) | Appropriate
121 | Rectum/Bladder/Anal cannal | Bolus | FALSE |  | Appropriate
122 | Rectum/Bladder/Anal cannal | Setup Notes | FALSE |  | Appropriate
123 | Rectum/Bladder/Anal cannal | Beam Energy | TRUE | 15 MV | Appropriate
124 | Rectum/Bladder/Anal cannal | Collimator Rotation | TRUE | 270 | Appropriate
125 | Rectum/Bladder/Anal cannal | Wedges | TRUE | EDW 45 (Rectum, Anal canal) | Appropriate
126 | Rectum/Bladder/Anal cannal | Patient & Beam Info | FALSE |  | Appropriate
127 | Rectum/Bladder/Anal cannal | MU_on & MU_off | FALSE |  | Appropriate
114 | Rectum/Bladder/Anal cannal | Slice Thickness | TRUE | 5 mm | Appropriate
129 | Rectum/Bladder/Anal cannal | Effective Depth | FALSE |  | Appropriate
118 | Rectum/Bladder/Anal cannal | Beam Arrangement | TRUE | 3 Field Lats & PA (Rectum, Anal canal), 4 fld Box (Bladder) | Appropriate
150 | Rectum/Bladder/Anal cannal | intent | TRUE | Curative Neoadjuvant Palliative | Appropriate
151 | Rectum/Bladder/Anal cannal | Lat. Tattoo Reference | TRUE | On Isocentre | Appropriate
152 | Rectum/Bladder/Anal cannal | Patient Anatomy | TRUE | Full Bladder | Appropriate
153 | Rectum/Bladder/Anal cannal | Organs at Risk | TRUE | Small Bowel, Total dose less than Bladder tolerance | Appropriate
154 | Rectum/Bladder/Anal cannal | Field Normalization Type | TRUE | 100% to iso | Appropriate
155 | Rectum/Bladder/Anal cannal | Isocentre Shifts | TRUE | no shifts | Appropriate
156 | Rectum/Bladder/Anal cannal | PTV and CTV coverage | TRUE |  | Appropriate
157 | Rectum/Bladder/Anal cannal | 180 E | FALSE |  | Appropriate
29483 | Rectum/Bladder/Anal cannal | Dose to Critical Organs | TRUE | Small Bowel: The volume receiving full dose must be < 2.0 l | Appropriate
29480 | Rectum/Bladder/Anal cannal | Treatment Orientation | TRUE | HFP | Appropriate
28964 | Rectum/Bladder/Anal cannal | Straightening tattoo | TRUE | 10 cm SUP of PA tattoo | Appropriate
28963 | Rectum/Bladder/Anal cannal | Image Registration | FALSE |  | Appropriate
128 | Rectum/Bladder/Anal cannal | Equivalent Path Length | FALSE |  | Appropriate
111 | Rectum/Bladder/Anal cannal | Patient Orientation | TRUE | Head First Prone | Appropriate
112 | Rectum/Bladder/Anal cannal | Immobilization | TRUE | Belly Board | Appropriate
110 | Rectum/Bladder/Anal cannal | Dose and Fractions | TRUE | 45/25, 5.4/3 or 9/5 (Rectum, Anal Canal) 52.5/20 (Bladder), | Appropriate
109 | Rectum/Bladder/Anal cannal | Site | TRUE | Pelvis, Rectum, Anal Canal, Bladder | Appropriate
29439 | Skin - Electrons | Lat. Tattoo Reference | FALSE |  | Appropriate
256 | Skin - Electrons | MLC | TRUE | None | Appropriate
255 | Skin - Electrons | Tolerance Table | TRUE | Electrons Planned | Appropriate
252 | Skin - Electrons | Wedges | FALSE |  | Appropriate
61 | Skin - Electrons | BEV Image Quality | FALSE |  | Appropriate
253 | Skin - Electrons | Portal Image Isocentre | FALSE |  | Appropriate
62 | Skin - Electrons | MLC Coverage | FALSE |  | Appropriate
69 | Skin - Electrons | Applicator Clearance | FALSE | Min 95 cm clearance | Appropriate
68 | Skin - Electrons | Smoothing | TRUE | 3D Gaussian Medium | Appropriate
67 | Skin - Electrons | Accuracy | TRUE | 1% | Appropriate
66 | Skin - Electrons | BBs & Wires Contoured out | TRUE |  | Correct
63 | Skin - Electrons | MLC Leaf Gap | FALSE |  | Appropriate
29153 | Skin - Electrons | Straightening tattoo | FALSE |  | Appropriate
254 | Skin - Electrons | 180 E | FALSE |  | Appropriate
194 | Skin - Electrons | Patient Anatomy | TRUE | Artifacts from Dental Work etc do not interfere with dose calculation. | Appropriate
53 | Skin - Electrons | Beam Arrangement | TRUE | Approximately skin apposition | Appropriate
45 | Skin - Electrons | Contrast | FALSE |  | Appropriate
46 | Skin - Electrons | Contrast | FALSE |  | Appropriate
47 | Skin - Electrons | Field Normalization Type | FALSE |  | Appropriate
29156 | Skin - Electrons | Immobilization | TRUE | Barker System | Appropriate
48 | Skin - Electrons | Heterogenety | FALSE |  | Appropriate
49 | Skin - Electrons | Grid Size | TRUE | 0.1 for 6&9 Mev, 0.15 for 12MeV, 0.2 for 16MeV, 0.25 for 20MeV | Appropriate
50 | Skin - Electrons | Calculation Model | TRUE | EMC_8117 | Appropriate
29155 | Skin - Electrons | Wedges | FALSE |  | Appropriate
52 | Skin - Electrons | Bolus Thickness and Extent | TRUE | Based on energy & tumour depth | Appropriate
44 | Skin - Electrons | Dose and Fractions | TRUE | 45/10, 55/20, or 66/33 | Appropriate
54 | Skin - Electrons | Isocentre Shifts | FALSE |  | Appropriate
55 | Skin - Electrons | SSD | TRUE | 100+ bolus thickness | Appropriate
56 | Skin - Electrons | Treatment Units The Same | FALSE |  | Appropriate
57 | Skin - Electrons | Collimator Rotation | FALSE |  | Appropriate
58 | Skin - Electrons | Wedges | FALSE |  | Appropriate
59 | Skin - Electrons | Field Jaw Settings | FALSE |  | Appropriate
60 | Skin - Electrons | Field Size | FALSE |  | Appropriate
51 | Skin - Electrons | GTV, CTV, PTV | TRUE | Contours based on wires and depth | Appropriate

## View
View_ID | View_Name | View_Document | View_Description | View_Sort_Order | View_show | Multi_Phase
1 | Prescription Sheet | Paper | RO Indicated Prescription & Technique | 1 | FALSE | FALSE
38 | Informed Concent | Paper | Patient Summary of Intent and complications | 2 | FALSE | FALSE
37 | Clinical Progress Notes | Paper | Clinical Patient Notes | 3 | FALSE | FALSE
5 | Technical Progress Notes | Paper | Notes regarding treatment planning Descisions | 4 | TRUE | FALSE
3 | CT Simulation Worksheet | Dynamic Documents |  | 5 | FALSE | FALSE
30 | Preliminary CT Plan Report | Dynamic Documents | CT Scan and setup parameters | 7 | TRUE | FALSE
31 | LapData Report | Dynamic Documents | LAP Laser Settings for tattooing in CT | 9 | FALSE | FALSE
55 | Course | Eclipse Selection | Course Properties | 9.5 | TRUE | FALSE
11 | CT Image Properites - Technical | Eclipse Selection |  | 10 | FALSE | FALSE
8 | CT Image Properites - CT Geometry | Eclipse Selection |  | 11 | FALSE | FALSE
33 | CT Volume Properties - 3D | Eclipse Selection | Slices used for volume contruction | 12 | FALSE | FALSE
9 | CT Volume Properties - Origin | Eclipse Selection | Plan Origin | 13 | TRUE | FALSE
10 | CT Volume Properties - User | Eclipse Selection |  | 14 | FALSE | FALSE
32 | CT Volume Orthogonal Views | Eclipse Selection | CT Volume Images | 15 | TRUE | FALSE
17 | Dose Prescription | Eclipse Field Setup |  | 15.3 | TRUE | TRUE
41 | Field Matching Worksheet | Paper | Spreadsheet for calculating Gantry Collimator and Couch angles for 3& 4 field breasts | 15.5 | TRUE | FALSE
19 | DVH | Eclipse Field Setup | Includes Dose Statistics Tab | 15.9 | TRUE | FALSE
54 | Optimize | Eclipse Field Setup | Optimization Calculations View | 15.95 | TRUE | FALSE
18 | Electron & Photon Calculation Models | Eclipse Field Setup | Selected under Calculation Models/ Options | 17 | TRUE | TRUE
49 | Optimization Calculation Models | Eclipse Field Setup | Selected under Calculation Models/ Options | 17.5 | TRUE | FALSE
13 | Axial Slices | Eclipse Field Setup | For Confirming location of BBs and Isocentre | 18 | TRUE | TRUE
14 | Field Parameters | Eclipse Field Setup | Tab on Bottom | 19 | TRUE | TRUE
40 | Initial Multi-Field Plan | Eclipse Field Setup | Contains individual fields to be merged for field-in-field | 21 | TRUE | FALSE
16 | Dose Distributions | Eclipse Field Setup |  | 22 | TRUE | TRUE
15 | BEVs | Eclipse Field Setup |  | 23 | TRUE | TRUE
20 | Reference Point Properties - Location | Eclipse Field Setup | For MUREF | 24 | TRUE | TRUE
29 | Eclipse and TPP Isocentre QA | Spreadsheets | For Verifying LapData and TPP Isocentre | 24.1 | TRUE | FALSE
42 | TPP Plan Data Sheet | Paper | Plan Datasheet Printout from TPP | 24.2 | TRUE | TRUE
43 | TPP Dose Distribution | TheraPlan Plus | Calculated Dose Distributions in TPP | 24.4 | TRUE | TRUE
44 | BEVs in TPP | TheraPlan Plus | BEVs in TPP | 24.6 | TRUE | TRUE
12 | Plan Properties - General | Eclipse Field Setup |  | 25 | TRUE | TRUE
35 | Dose Distribution | Dynamic Documents | Screen Shot of axial distribution | 26 | TRUE | TRUE
36 | Final Plan Report | Dynamic Documents | Screen Shot of Final Plan Report | 27 | TRUE | TRUE
21 | Final Plan Report | Eclipse Field Setup | Preview in Eclipse | 27.5 | TRUE | TRUE
39 | IMSure User Factor | Spreadsheets | For Determining an Inhomogeneity User Factor for IMSURE | 29 | TRUE | TRUE
28 | IMSure Calc Points | IMSure | Dose Calculation Points for IMSure | 30 | TRUE | TRUE
26 | IMSure Photons & Electrons | IMSure | Photon MU Calculations | 31 | TRUE | TRUE
45 | MU Calc | Spreadsheets | Second MU calculations for TPP | 32.5 | TRUE | TRUE
47 | Portal Image Plans | Eclipse Field Setup | Image only plans. No wedges | 32.7 | TRUE | TRUE
22 | RT Chart Prescription | RT Chart | Plan Prescriptions in RT Chart | 33 | TRUE | TRUE
56 | Scheduling | RT Chart | Fraction & Imaging Schedule | 33.5 | TRUE | TRUE
23 | RT Chart Field Parameters | RT Chart | Final Field Parameters | 34 | TRUE | TRUE
25 | RT Chart Reference Points | RT Chart | Dose Limits and Treatment Approval | 35 | TRUE | TRUE
52 | EPID Dose QA Report | Dynamic Documents | Record of EPID IMRT QA | 35.1 | TRUE | FALSE
53 | Matrixx Dose QA Report | Dynamic Documents | Record of Matrix IMRT QA | 35.2 | TRUE | FALSE
2 | QA Task | Time Planner | Plan Process tracking | 36 | TRUE | FALSE
48 | QA task sheet | Paper | Paper QA record | 37 | TRUE | FALSE
50 | Activity Capture | Time Planner | Radiotherapy treatment record | 38 | TRUE | FALSE
51 | CRTP | OPIS | Radiation Treatment Plan Track | 39 | FALSE | FALSE
57 | IMRT EPID Spreadsheet | Spreadsheets | Record of IMRT EPID QA for Tracking Purposes | 40 | TRUE | FALSE

## Document
Document_ID | Document_Name | Document_Description | Document_Sort_Order
1 | Paper | RO Indicated Prescription & Technique | 1
7 | Dynamic Documents | Setup & Tattoo Reference for CT Scan | 2
20 | Spreadsheets | For Verifying LapData and TPP Isocentre | 3
11 | Eclipse Selection | CT and Volume Data in Eclipse | 4
12 | Eclipse Field Setup | Plan and Field Data in Eclipse | 5
17 | IMSure | Photon MU Calculations | 6
13 | RT Chart | Plan Prescriptions in RT Chart | 7
2 | Time Planner | Plan Process tracking | 8
22 | TheraPlan Plus | Theraplan Plus Electronic Records | 9
23 | OPIS | OPIS Electronic Activity Capture | 10

## Category
Category_ID | Category_Name | Category_Description | Category_Sort_Order
1 | Patient ID | Identifies the Patient | 1
2 | Plan ID | Identifies the Plan | 1.5
5 | Scan Data | CT Scan Parameters | 2
9 | Treatment Technique | Type of plan e.g 4-field Box, Wedged pair, Bolus | 2
3 | Prescription | Dose Prescription | 3
20 | Image Quality | BEV / CT Image Quality & Annotation | 3
6 | Set-up | Patient Setup Information | 4
12 | Contours | Presence and quality of contours drawn | 4.5
17 | Isocentre | Parameters identifying the Isocentre location including origin, shifts | 5
15 | Calculation Parameters | Algorithm, grid setting, inhomogeneity | 6
8 | Field Parameters | Field Size, Gantry Angle | 7
18 | Dose data | Distribution uniformity, hot spots, PTV, CTV dose | 8
16 | Comments | Non-required information | 9
13 | Document Validated | Signatures of electronic approval | 10

## Criteria
Criteria_ID | Criteria_Name | Criteria_Description
1 | Appropriate | The Item has an expected value or range of values
2 | Present | The item's value must be present but is not evaluated
3 | Consistent | The items value agrees with the value on another view
4 | Entered | Data Enetered While Checking
5 | Correct | Item is true
6 | For Information | Item is not checked, but information is obtained from it

## Disease
Disease_ID | Disease_Name | Disease_Description | Disease_Sort_Order | Default_number_of_Phases
6 | Generic | All Other Disease Sites | 0 | 2
19 | Generic IMRT | IMRT for All Other Disease Sites | 0.1 | 1
1 | Breast Tangents | Breast Treated with tangent pair only (Field-in-field) | 1 | 1
2 | Breast with Nodes | 3 & 4 Field Breast Technique | 2 | 2
13 | IMRT Prostate | Low, Post Op & Intermediate Risk Prostate using IMRT | 3 | 1
18 | Prostate 2 Phase with IMRT | High Risk Prostate with Hormones using IMRT for Phase II | 3.5 | 2
3 | Prostate Single Phase | Low, Post Op & Intermediate Risk Prostate | 3.7 | 1
9 | Prostate 2 Phase | High Risk Prostate with Hormones | 4 | 2
5 | Lung |  | 5 | 1
7 | Skin - Electrons | Skin Electron treatments | 6 | 1
11 | IMRT Head & Neck | Head & Neck Planned with IMRT | 8 | 1
15 | IMRT CNS | Brain IMRT | 10 | 1
12 | GYN | Gynae and Pelvis region | 11 | 1
4 | Rectum/Bladder/Anal cannal | Cancers of the Recum Anal canal and Bladder | 11.5 | 2
16 | Pancreas |  | 11.7 | 1
14 | Esophagus | Conventionally Planned Esophagus | 12 | 1
17 | Brain | Glioma, PCI, Conventional brain | 13 | 1
10 | H&N w Comps | Head & Neck Planned in TPP with Compensators | 100 | 3